## Key to Data Base Fields-Columns
Entry # | Priority | Status / Action | Area | Date | Time | run # (During or Before) | Source | TS / Pcrate ID | Pcrate slot | chamber | On-chamber board | Problem Description | Confidence of Problem Solution | Estimated Repair Time | Comments on Action Needed | Chamber Summary | Check VME | check ME
Example
13 | 1 | Study |  | 2008-06-18 |  |  | Jason | VMEp1_1 |  | ME+1/1/01 | CFEB5 | dead/flakey comparators |  | 5h | ME+1/1 removal
db entry number
Priority to be applied to corrective action
|  | 5 |  | Excluded from GLOBAL running
|  | 4 |  | Entire chamber analog or digital not working > ALCT dead, 3+ CFEBs dead, HV for chamber or plane dead
|  | 3 |  | Partial chamber not working > 1-3 CFEBS, HV sector
|  | 2 |  | Usually minor problems
|  | 1 |  | Occasional minor problems
|  | 0 |  | Problems fixed / working OK
Required Action
|  |  | remove N |  | remove N chambers to fix (including chamber itself)
|  |  | acess S+(-)m |  | requires special access Station +(-)m
|  |  | study |  | requires more study before taking remedial action
|  |  | other |  | other actions
Work Sheets
|  | Summary |  |  |  | Lists chambers, one line per chamber, sorted by priotity
copied from the source page and re-sorted
|  | Source - Sort by CSC, Date |  |  |  | Source | One row per entry per chamber
|  | … |  |  |  |  | each entry is one instance of bad behavior for the chamber
this is primarily a log of chamber problems seen

## Summary
Entry # | Priority | Chamber Action | Area | Last Update Date | Time | run # (During or Before) | Source | TS / Pcrate ID | Pcrate slot | chamber | On-chamber board | Problem Description | Confidence of Solution | Estimated Repair Time | Comments on Action Needed | Chamber Summary | Check VME | check ME
516 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_7 |  | ME+1/1/19 | ALCT: | Config problems for ALCT: ME+1/1/19, known threshold problem due to dead delay chip + corrupted data | 90 | 5h | remove chamber | Y
517 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_7 |  | ME+1/1/20 | ALCT | Replace ALCT | 90 | 5h | remove chamber | Y
518 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_8 |  | ME+1/1/22 | ALCT | no ALCT trigger, replace ALCT | 90 | 5h | ME+1/1 removal | Y
519 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_9 |  | ME+1/1/26 | CFEB | CFEB2 & CFEB5 need to be replaced | 90 | 5h | remove chamber | Y
523 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_6 |  | ME+1/2/15 | ALCT | bad fuse on LVDB replace | 90 | 8h | remove lvdb >> 2+ CSCs | Y
525 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp2_1 |  | ME+2/1/03 | ALCT | Config problems for ALCT most likely 1.8 V fuse blown, replace ALCT | 90 | 8h | back, remove 1-2 CSC, replace ALCT | Y
550 | 4 |  |  | 2008-07-20 |  |  | Misha | VMEp3_3 |  | ME+3/1/10 | ALCT | hot (trigger pass always "ON") |  |  |  | Y
530 | 4 | study |  | 2008-07-21 |  |  | Misha | VMEp3_6 |  | ME+3/1/17 | ALCT | data format problem: ALCT Not Present - present 18-July run | 0.2 | 2h | remove 2+ CSCs. | Y
532 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEm1_2 |  | ME-1/1/04 | ALCT | Fuse out | 100 | 8h | pull chamber | Y
537 | 4 | study |  | 2008-07-18 |  |  | Misha | VMEm1_11 |  | ME-1/3/31 | ALCT | ALCT was swapped, but problem remains, problem unknown | 100 | 2h | ALCT swap | Y
17 | 4 | access |  | 2008-07-18 |  |  | Misha | VMEm2_6 |  | ME-2/2/36 | ALCT | TMB slot 16 ALCT 1.8 V fuse blown. | 90 | 5h | need YE-2 open | Y
542 | 4 | leave |  | 2008-07-18 |  |  | Misha | VMEm3_1 |  | ME-3/2/03 | HV | Turn off HV sector + disable interlock | 90 | 0h | external action only | Y
207 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEm3_4 |  | ME-3/2/23 | ALCT | always sends corrupted data - have to kill ALCT | 90 | 8h | remove front chambers | Y
515 | 3 | remove |  | 2008-07-18 |  |  | Jason | VMEp1_4 |  | ME+1/1/11 | CFEB | CFEB#4 comparator Problem | 90 | 5h | ME+1/1 removal | Y
553 | 3 |  |  | 2008-07-21 |  |  | Greg | VMEp1_12 |  | ME+1/1/35 | CFEB1 | kill_input=3 |  |  |  | Y
521 | 3 | study |  | 2008-07-18 |  |  | Jason | VMEp1_1 |  | ME+1/2/01 | CFEB3 | BAD CFEB3/DMB, was excluded from CRUZET3, replace CFEB | 20 | 2h | access to +1 | Y
522 | 3 | access |  | 2008-07-18 |  |  | Jason | VMEp1_5 |  | ME+1/2/13 | CFEB5 | DMB-CFEB LCT mismatches | 90 | 8h | remove 2+ CSCs | Y
547 | 3 |  |  | 2008-07-20 |  |  | Misha | VMEp1_8 |  | ME+1/2/21 | CFEB1 & CFEB5 | comparator DAC has higher offest for CFEB1, working fine. Killed CFEB5 |  |  |  | Y
545 | 3 |  |  | 2008-07-20 |  |  | Misha | VMEp2_3 |  | ME+2/2/15 | CFEB5 | 80MHz mux clock dead, the SCA data path OK |  |  |  | Y
529 | 3 | remove |  | 2008-07-18 |  |  | Jason | VMEp3_6 |  | ME+3/1/17 | CFEB1 | dead comps. & SCA, configure problem - under investigation - ALCT? | 20% | 8h | remove 2+ CSCs. | Y
554 | 3 |  |  | 2008-07-21 |  |  | Greg | VMEp3_6 |  | ME+3/2/35 | CFEB4 | kill_input=6 |  |  |  | Y
533 | 3 | remove |  | 2008-07-18 |  |  | Misha | VMEm1_9 |  | ME-1/1/25 | HV | HV channel will not stay at full voltage - 2800 | 90 | 8h | disable HV channel - replace chamber | Y
534 | 3 | study |  | 2008-07-09 |  |  | Jason | VMEm1_12 |  | ME-1/1/35 | CFEB4 | Bad CFEB4 needs more study | 20 | 5h |  | Y
538 | 3 | remove 3 |  | 2008-07-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/03 | CFEB2 | dead - must swap CFEB | 100 | 8h | swap >access, remove front chambers | Y
204 | 3 | remove |  | 2008-07-18 |  |  | Misha | VMEm3_4 |  | ME-3/1/13 | CFEB3 | dead | 100 | 5h | swap >access, remove front chambers | Y
205 | 3 | remove |  | 2008-07-18 |  |  | Misha | VMEm3_5 |  | ME-3/1/15 | CFEB2 | dead channel link serial bit (middle bit) is broken | 90 | 5h | swap >access, remove front chambers | Y
206 | 3 | access |  | 2008-07-02 |  |  | Misha | VMEm3_4 |  | ME-3/2/22 | CFEB3 | dead channel link serial bit (middle bit) is broken | 90 | 2h | swap >access | Y
554 | 2 | study |  | 2008-07-20 |  |  | Misha | VMEp1_8 |  | ME+1/1/23 | CFEB3 | under investigation |  |  |  | Y
526 | 2 | study |  | 2008-07-18 |  |  | Misha | VMEp2_2 |  | ME+2/2/14 | ALCT | Sometimes ALCT bad, replace ALCT | 90 | 2h | access | Y
528 | 2 | access |  | 2008-07-18 |  |  | Misha | VMEp3_1 |  | ME+3/1/04 | CFEB | Hot comparator - check cables at chamber - exclude CFEB until fixed | 90 | 2h | access to chamber | Y
555 | 2 | study |  | 2008-07-20 |  |  | Misha | VMEp3_4 |  | ME+3/2/21 | CFEB3 | bad CRCs |  |  |  | Y
531 | 2 | remove |  | 2008-07-18 |  |  | Jason | VMEm1_2 |  | ME-1/1/03 | CFEB5 | hot comparator pl 1,4 - one channel hot only | 80 | 8h | pull chamber | Y
540 | 2 | remove |  | 2008-07-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/06 | HV | one HV sector constant trip; Alex states the problem in in the chamber | 90 | 8h | replace chamber | Y
541 | 2 | access |  | 2008-07-18 |  |  | Misha | VMEm3_4 |  | ME-3/1/12 | CFEB1 | Hot SCA channels layer 1 CFEB1 - always there | 90 | 2h | access chamber and swap CFEB | Y
548 | 1 |  |  | 2008-07-20 |  |  | Jianhui | VMEp2_3 |  | ME+2/1/10 | CFEB4 | comparator DAC has higher offest for CFEB1, working fine. |  |  |  | Y
543 | 1 |  |  | 2008-07-20 |  |  | Jianhui | VMEm1_10 |  | ME-1/1/29 | CFEB5 | JTAG TDO read problem |  |  |  | Y

## Source - Sort by CSC, VME, DATE
Entry # | Priority | Chamber Action | Area | Date | Time | run # (During or Before) | Source | TS / Pcrate ID | Pcrate slot | chamber | On-chamber board | Problem Description | Confidence of Problem Solution | Estimated Repair Time | Comments on Action Needed | Chamber Summary | Check VME | check ME
13 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_1 |  | ME+1/1/01 | CFEB5 | dead/flakey comparators
221 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_1 |  | ME+1/1/01 | CFEB5 | dead/flakey comparators
398 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_1 |  | ME+1/1/02 |  | ~1% CFEB CRC errors (DMB was replaced)
63 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEp1_1 |  | ME+1/1/02 | CFEB | - ~1% CFEB CRC errors (minor, but this is in every run)
145 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_1 |  | ME+1/1/02 |  | ~1
619 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/1/06 |  | 5 cfeb3 - try reprogram failed
515 | 3 | remove |  | 2008-07-18 |  |  | Jason | VMEp1_4 |  | ME+1/1/11 | CFEB | CFEB#4 comparator Problem | 90 | 5h | ME+1/1 removal | Y
455 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_4 |  | ME+1/1/11 |  | CFEB#4 DAV-LCT match errors, sometimes accompanied by a Full FIFO@DMB.
400 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_4 |  | ME+1/1/11 |  | 85% CFEB4 FIFO Full, 20% B-Words, CFEB4 L1A is out of sync, ~1% CFEB DAV, Sample Count errors, ALCT L1A out of sync (constant offset -6 counts) [high]
11 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_4 |  | ME+1/1/11 | CFEB4 | comparator problem
219 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_4 |  | ME+1/1/11 | CFEB4 | comparator problem
146 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_4 |  | ME+1/1/11 |  | 85 B-Words, CFEB4 L1A is out of sync, ~1 CFEB DAV, Sample Count errors,
576 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_5 |  | ME+1/1/13 | ALCT: | Config problems for
504 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_5 |  | ME+1/1/13 |  | Config problems for ALCT AFEB 17 wrong reading of threshold 241/68(expected)
43 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_5 |  | ME+1/1/13 | ALCT: | Config problems for ALCT: ME+1/1/13, Check TMB+ALCT config make problem go away
620 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/1/19 |  | kill at DDU timing checked to be OK
516 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_7 |  | ME+1/1/19 | ALCT: | Config problems for ALCT: ME+1/1/19, known threshold problem due to dead delay chip + corrupted data | 90 | 5h | remove chamber | Y
507 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_7 |  | ME+1/1/19 |  | Config problems for ALCT
457 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_7 |  | ME+1/1/19 |  | Bad ALCT/TMB
348 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_7 |  | ME+1/1/19 | ALCT: | Config problems for
314 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_7 |  | ME+1/1/19 | ALCT: | Config problems for
298 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_7 |  | ME+1/1/19 | ALCT
262 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_7 |  | ME+1/1/19 | ALCT | Config problems for ALCT
414 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_7 |  | ME+1/1/19 |  | (was masked) ~100% ALCT FIFO Full and DMB-Input timeouts for ALCT, ALCT L1A out of sync - messed up
106 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_7 |  | ME+1/1/19 | ALCT | Config problems for ALCT: ME+1/1/19,
45 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_7 |  | ME+1/1/19 | ALCT: | Config problems for ALCT: ME+1/1/19, known threshold problem due to dead delay chip
14 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_7 |  | ME+1/1/19 | ALCT | configure problem |  | 5h | ME+1/1 removal
222 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_7 |  | ME+1/1/19 | ALCT | configure problem | 20% | 5h | ME+1/1 removal. Next replace TMB/RAT
326 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_7 |  | ME+1/1/19 | ALCT | configure problem | 0.2 | 5h | ME+1/1 removal. Next replace TMB/RAT
147 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_7 |  | ME+1/1/19 |  | (was masked) ~100 messed up
577 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_7 |  | ME+1/1/19, | ALCT: | Config problems for
476 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_7 |  | ME+1/1/19, |  | Config problems for ALCT:
621 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/1/20 |  | kill at DDU timing checked to be OK
517 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_7 |  | ME+1/1/20 | ALCT | Replace ALCT | 90 | 5h | remove chamber | Y
456 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_7 |  | ME+1/1/20 |  | Bad ALCT/RAT
415 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_7 |  | ME+1/1/20 |  | (was masked) 72% ALCT CRC Errors
1 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_7 |  | ME+1/1/20 | ALCT | (data format problem)
210 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_7 |  | ME+1/1/20 | ALCT | problem:bad data consistency, CRCs seen at TMB | 20% | 5h | remove chamber. Replaced TMB 06.09.08 10:00. replace RAT next
148 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_7 |  | ME+1/1/20 |  | (was masked) 72
622 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/1/21 |  | kill at DDU added to kill list today
623 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/1/22 |  | 1 may need to swap TMB or RAT - timing checked to be OK
559 |  |  |  | 2008-07-22 |  |  | Luca | VMEp1_8 |  | ME+1/1/22 | DMB | chip 1 1.8 V reads as 1844 mV (also problems with ALCT config sometimes)
558 |  |  |  | 2008-07-21 |  |  | Greg | VMEp1_8 |  | ME+1/1/22 | ALCT | kill_input=1
618 |  |  |  | 2008-07-20 |  |  | Misha | VMEp1_8 |  | ME+1/1/22 | ALCT d | ead in terms of capability of senting out triggers
518 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_8 |  | ME+1/1/22 | ALCT | no ALCT trigger, replace ALCT | 90 | 5h | ME+1/1 removal | Y
371 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEp1_8 |  | ME+1/1/22 | RUI#11, DDU Input 0: | ALCT L1A Error set 8 times, Consistency:9, CRC:9
384 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEp1_8 |  | ME+1/1/22 | RUI#11, DDU Input 0: | ALCT CRC Error set 141 times, Consistency:141
404 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_8 |  | ME+1/1/22 |  | 76% ALCT CRC and ALCT word count errors [high]
64 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEp1_8 |  | ME+1/1/22 | ALCT | - 81% ALCT CRC and ALCT word count errors
87 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEp1_8 |  | ME+1/1/22 | CFEB | - “dead”; CFEB problems? (wires look fine)
12 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_8 |  | ME+1/1/22 | ALCT | problem
220 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_8 |  | ME+1/1/22 | ALCT | problem: bad data consistency, CRCs seen at TMB | 20% | 5h | ME+1/1 removal. Next replace TMB/RAT
149 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_8 |  | ME+1/1/22 |  | 76
578 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_8 |  | ME+1/1/23 | DMB: | Config problems for
554 | 2 | study |  | 2008-07-20 |  |  | Misha | VMEp1_8 |  | ME+1/1/23 | CFEB3 | under investigation |  |  |  | Y
616 |  |  |  | 2008-07-20 |  |  | Misha | VMEp1_8 |  | ME+1/1/23 | CFEB#2 | JTAG TDO read problem.
508 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_8 |  | ME+1/1/23 | CFEB 1 | CFEB firmware incorrect
509 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_8 |  | ME+1/1/23 |  | Config problems for DMB ME+1/2/21
349 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_8 |  | ME+1/1/23 | DMB: | Config problems for
315 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_8 |  | ME+1/1/23 | CFEB | firmware incorrect:
316 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_8 |  | ME+1/1/23 | DMB: | Config problems for
300 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEp1_8 |  | ME+1/1/23 | CFEB 1
264 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_8 |  | ME+1/1/23 | DMB | Config problems for DMB
281 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEp1_8 |  | ME+1/1/23 |  | CFEB firmware incorrect
107 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_8 |  | ME+1/1/23 | DMB | Config problems for DMB:
46 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_8 |  | ME+1/1/23 | DMB: | Config problems for DMB: ME+1/1/23, ME+1/2/21, well known problem on both of them. One is problem with temp
571 |  |  |  | 2008-07-22 |  |  | jason | VMEp1_8 |  | ME+1/1/23, |  | replace CFEB2 (Gu, lower priority)
477 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_8 |  | ME+1/1/23, |  | Config problems for DMB:
626 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/1/26 |  | kill at DDU cfeb2,3?,4-thresholds, 5
556 |  |  |  | 2008-07-21 |  |  | Greg | VMEp1_9 |  | ME+1/1/26 | CFEB2 | kill_input=4
613 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+1/1/26 | CFEB#2 | FPGA not programmed |  |  |  |  | XXX
614 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+1/1/26 | CFEB#5 | dead |  |  |  |  | XXX
519 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_9 |  | ME+1/1/26 | CFEB | CFEB2 & CFEB5 need to be replaced | 90 | 5h | remove chamber | Y
510 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_9 |  | ME+1/1/26 |  | Config problems for DMB
458 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_9 |  | ME+1/1/26 |  | DMB FIFO Full
351 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_9 |  | ME+1/1/26 | DMB: | Config problems for
317 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_9 |  | ME+1/1/26 | DMB: | Config problems for
302 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_9 |  | ME+1/1/26 | DMB
266 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_9 |  | ME+1/1/26 | DMB | Config problems for DMB
401 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_9 |  | ME+1/1/26 |  | ~100% CFEB5 DMB-FIFO full (starting from 080624_090805), ~30% CFEB DAV errors, CFEB2 is masked [high]
108 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_9 |  | ME+1/1/26 | DMB | Config problems for DMB: ME+1/1/26,
48 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_9 |  | ME+1/1/26 | DMB: | Config problems for DMB: ME+1/1/26, well known problem. Went to DMB tests, fails DMB test 8: CFEB4 comparator threshold reading wrong
2 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_9 |  | ME+1/1/26 | CFEB5 | (data format, dead comps. & SCA)
211 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_9 |  | ME+1/1/26 | CFEB5 | data format, dead comps. & SCA, config problem | 90% | 5h | remove chamber. Replaced DMB 21.04.08 10:00. still bad
150 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_9 |  | ME+1/1/26 |  | ~100 CFEB DAV errors, CFEB2 is masked
580 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_9 |  | ME+1/1/26, | DMB: | Config problems for
479 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_9 |  | ME+1/1/26, |  | Config problems for DMB:
291 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_11 |  | ME+1/1/32 | DMB
259 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_11 |  | ME+1/1/32 | DMB | Config problems for DMB
277 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEp1_11 |  | ME+1/1/32 |  | CFEB firmware incorrect
520 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEp1_12 |  | ME+1/1/34 | ALCT | Now OK, need to watch - had blown fuse type symtoms |  |  |  | Y
500 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_12 |  | ME+1/1/34 |  | ALCT firmware incorrect
502 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_12 |  | ME+1/1/34 |  | Config problems for ALCT most likely 1.8 V fuse blown
553 | 3 |  |  | 2008-07-21 |  |  | Greg | VMEp1_12 |  | ME+1/1/35 | CFEB1 | kill_input=3 |  |  |  | Y
609 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+1/1/35 | CFEB#1 | dead |  |  |  |  | XXX
501 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_12 |  | ME+1/1/35 | CFEB 0 | CFEB firmware incorrect
503 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_12 |  | ME+1/1/35 |  | Config problems for DMB
352 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_12 |  | ME+1/1/35 | DMB: | Config problems for
309 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_12 |  | ME+1/1/35 | CFEB | firmware incorrect:
310 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_12 |  | ME+1/1/35 | DMB: | Config problems for
292 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_12 |  | ME+1/1/35 | DMB
293 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEp1_12 |  | ME+1/1/35 | CFEB 0
260 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_12 |  | ME+1/1/35 | DMB | Config problems for DMB
278 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEp1_12 |  | ME+1/1/35 |  | CFEB firmware incorrect
416 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_12 |  | ME+1/1/35 |  | no comparators data for CFEB1
109 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_12 |  | ME+1/1/35 | DMB | Config problems for DMB: ME+1/1/35,
49 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_12 |  | ME+1/1/35 | DMB: | Config problems for DMB: ME+1/1/35, no communication with CFEB1
151 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_12 |  | ME+1/1/35 |  | ME+1/1/35 – no comparators data for CFEB3
572 |  |  |  | 2008-07-22 |  |  | jason | VMEp1_12 |  | ME+1/1/35, |  | replace CFEB1 (Gu)
581 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_12 |  | ME+1/1/35, | DMB: | Config problems for
474 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_12 |  | ME+1/1/35, |  | Config problems for DMB:
627 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/2/01 |  | 5 cfeb3
551 |  |  |  | 2008-07-21 |  |  | Greg | VMEp1_1 |  | ME+1/2/01 | CFEB3 | kill_input=5
521 | 3 | study |  | 2008-07-18 |  |  | Jason | VMEp1_1 |  | ME+1/2/01 | CFEB3 | BAD CFEB3/DMB, was excluded from CRUZET3, replace CFEB | 20 | 2h | access to +1 | Y
499 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_1 |  | ME+1/2/01 |  | Config problems for DMB
459 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_1 |  | ME+1/2/01 |  | BAD CFEB3/DMB
345 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_1 |  | ME+1/2/01 | DMB: | Config problems for
308 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_1 |  | ME+1/2/01 | DMB: | Config problems for
290 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_1 |  | ME+1/2/01 | DMB
258 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_1 |  | ME+1/2/01 | DMB | Config problems for DMB
412 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_1 |  | ME+1/2/01 |  | (was masked) CFEB3 is masked (no SCA and comparators data)
110 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_1 |  | ME+1/2/01 | DMB | Config problems for DMB: ME+1/2/1,
41 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_1 |  | ME+1/2/01 | DMB: | Config problems for DMB: ME+1/2/1, failed DMB test 8
5 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_1 |  | ME+1/2/01 | CFEB3 | (config problem) |  | 8h | remove 2+ CSCs
142 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_1 |  | ME+1/2/01 |  | (was masked) CFEB3 is masked (no SCA and comparators data)
575 |  |  |  | 2008-07-22 |  |  | jason | VMEp1_1 |  | ME+1/2/01, |  | replace CFEB3 (study has shown it is bad)
473 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_1 |  | ME+1/2/01, |  | Config problems for DMB:
42 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_2 |  | ME+1/2/04 | ALCT: | Config problems for ALCT: ME+1/2/4, Voltage reads too high
582 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_1 |  | ME+1/2/1 | DMB: | Config problems for
607 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+1/2/1 | CFEB#3 | FPGA not programmed |  |  |  |  | XXX
346 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_5 |  | ME+1/2/12 | DMB: | Config problems for
311 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_5 |  | ME+1/2/12 | DMB: | Config problems for
294 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_5 |  | ME+1/2/12 | DMB
295 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEp1_5 |  | ME+1/2/12 | DMB Control FPGA
279 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEp1_5 |  | ME+1/2/12 |  | DMB Control FPGA firmware incorrect
641 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/2/13 |  | remeasured timings, changed tmb_lct_cable_delay from 3 -> 2
522 | 3 | access |  | 2008-07-18 |  |  | Jason | VMEp1_5 |  | ME+1/2/13 | CFEB5 | DMB-CFEB LCT mismatches | 90 | 8h | remove 2+ CSCs | Y
460 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_5 |  | ME+1/2/13 |  | DMB-CFEB LCT mismatches
9 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_5 |  | ME+1/2/13 | CFEB5 | (DAV-LCT mismatch)
217 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_5 |  | ME+1/2/13 | CFEB5 | (DAV-LCT mismatch) | 20% | 8h | remove 2+ CSCs. Next replace DMB
628 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/2/15 |  | 1 LVDB blown fuse 3.3V ALC
550 |  |  |  | 2008-07-21 |  |  | Greg | VMEp1_6 |  | ME+1/2/15 | ALCT | kill_input=1
523 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp1_6 |  | ME+1/2/15 | ALCT | bad fuse on LVDB replace | 90 | 8h | remove lvdb >> 2+ CSCs | Y
505 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_6 |  | ME+1/2/15 |  | ALCT firmware incorrect
506 |  |  |  | 2008-07-17 |  |  | CC | VMEp1_6 |  | ME+1/2/15 |  | Config problems for ALCT
461 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp1_6 |  | ME+1/2/15 |  | Bad ALCT
347 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_6 |  | ME+1/2/15 | ALCT: | Config problems for
312 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_6 |  | ME+1/2/15 | ALCT | firmware incorrect:
313 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp1_6 |  | ME+1/2/15 | ALCT: | Config problems for
296 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_6 |  | ME+1/2/15 | ALCT
297 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEp1_6 |  | ME+1/2/15 | ALCT
261 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_6 |  | ME+1/2/15 | ALCT | Config problems for ALCT
280 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEp1_6 |  | ME+1/2/15 |  | ALCT firmware incorrect
413 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_6 |  | ME+1/2/15 |  | (was masked) HOT, 100% ALCT DAV Errors, 100% DMB-Input timeouts for ALCT
111 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_6 |  | ME+1/2/15 |  | Config problems for ALCT: ME+1/2/15,
44 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp1_6 |  | ME+1/2/15 | ALCT: | Config problems for ALCT: ME+1/2/15,
3 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_6 |  | ME+1/2/15 | ALCT | (bad fuse, config problem) |  | 8h | remove lvdb >> 2+ CSCs
143 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_6 |  | ME+1/2/15 |  | (was masked) HOT, 100 DMB-Input timeouts for ALCT
583 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_6 |  | ME+1/2/15, | ALCT: | Config problems for
475 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_6 |  | ME+1/2/15, |  | Config problems for ALCT:
320 |  |  |  | 2008-06-18 |  |  | Jason | VMEp1_6 |  | ME+1/2/16 | CFEB2 | (data error) |  | done | cfeb replaced
524 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEp1_7 |  | ME+1/2/20 | CFEB | Now OK, need to watch |  |  |  | Y
306 |  |  |  | 2008-07-03 |  |  | Misha | VMEp1_7 |  | ME+1/2/20 | CFEB | Needs to be times in
299 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_7 |  | ME+1/2/20 | DMB
263 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_7 |  | ME+1/2/20 | DMB | Config problems for DMB
403 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_7 |  | ME+1/2/20 |  | very low CFEB SCA data efficiency for all 5 CFEBs [low]
112 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_7 |  | ME+1/2/20 | DMB | Config problems for DMB: ME+1/2/20,
65 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEp1_7 |  | ME+1/2/20 | CFEB | - very low CFEB SCA data efficiency for all 5 CFEBs
88 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEp1_7 |  | ME+1/2/20 | DFEB | - “bad”; CFEB problems? (wires look fine)
144 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_7 |  | ME+1/2/20 |  | very low CFEB SCA data efficiency for all 5 CFEBs
624 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+1/2/21 |  | kill at DDU from Misha - timing checked to be OK
555 | 3 |  |  | 2008-07-21 |  |  | Greg | VMEp1_8 |  | ME+1/2/21 | CFEB5 | kill_input=7
547 | 3 |  |  | 2008-07-20 |  |  | Misha | VMEp1_8 |  | ME+1/2/21 | CFEB1 & CFEB5 | comparator DAC has higher offest for CFEB1, working fine. Killed CFEB5 |  |  |  | Y
612 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+1/2/21 | CFEB#1 | comparator DAC has higher offset, working fine |  |  |  |  | XXX
350 |  |  |  | 2008-07-06 |  |  | CC | VMEp1_8 |  | ME+1/2/21 | DMB: | Config problems for
301 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp1_8 |  | ME+1/2/21 | DMB
265 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp1_8 |  | ME+1/2/21 | DMB | Config problems for DMB
113 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp1_8 |  | ME+1/2/21 | DMB | Config problems for DMB:
47 | [6 empty] | Check Config | VMEp1_8 |  | ME+1/2/21 | DMB: | reading, one with minor voltage problem
579 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_8 |  | ME+1/2/21, | DMB: | Config problems for
478 |  |  |  | 2008-07-09 |  |  | CC | VMEp1_8 |  | ME+1/2/21, |  | Config problems for DMB:
606 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+1/2/25 | ALCT d | ead ---> ME+1/2/15 ALCT dead due to blown fuse on the LVDB |  |  |  |  | XXX
399 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp1_2 |  | ME+1/3/04 |  | higher than average SCA occupancies for CFEB3 (see “other issues 5” plots) [low]
141 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp1_2 |  | ME+1/3/04 |  | ME+1/3/04 – higher than average SCA occupancies for CFEB3 (see “other issues 5” plots)
584 |  |  |  | 2008-07-22 |  |  | CC | VMEp1_5 |  | ME+1/3/12 | ALCT: | Config problems for
410 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp2_6 |  | ME+2/1/01 |  | hot SCA channel CFEB3 layer 3 (old)
139 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp2_6 |  | ME+2/1/01 |  | ME+2/1/01 – hot SCA channel CFEB3 layer 3 (old)
557 | 4 |  |  | 2008-07-21 |  |  | Greg | VMEp2_1 |  | ME+2/1/03 | ALCT | kill_input=1
617 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+2/1/03 | ALCT " | dead" most likely 1.8 V fuse blown (still there is a |  |  |  |  | XXX
525 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEp2_1 |  | ME+2/1/03 | ALCT | Config problems for ALCT most likely 1.8 V fuse blown, replace ALCT | 90 | 8h | back, remove 1-2 CSC, replace ALCT | Y
511 |  |  |  | 2008-07-17 |  |  | CC | VMEp2_1 |  | ME+2/1/03 |  | ALCT firmware incorrect
512 |  |  |  | 2008-07-17 |  |  | CC | VMEp2_1 |  | ME+2/1/03 |  | Config problems for ALCT most likely 1.8 V fuse blown
548 | 1 |  |  | 2008-07-20 |  |  | Jianhui | VMEp2_3 |  | ME+2/1/10 | CFEB4 | comparator DAC has higher offest for CFEB1, working fine. |  |  |  | Y
615 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+2/1/10 | CFEB#4 | Comparator DAC has higher offset, working fine |  |  |  |  | XXX
513 |  |  |  | 2008-07-17 |  |  | CC | VMEp2_3 |  | ME+2/1/10 |  | Config problems for DMB
353 |  |  |  | 2008-07-06 |  |  | CC | VMEp2_3 |  | ME+2/1/10 | DMB: | Config problems for
318 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp2_3 |  | ME+2/1/10 | DMB: | Config problems for
303 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp2_3 |  | ME+2/1/10 | DMB
267 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp2_3 |  | ME+2/1/10 | DMB | Config problems for DMB
114 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp2_3 |  | ME+2/1/10 | DMB | Config problems for DMB: ME+2/1/10,
50 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEp2_3 |  | ME+2/1/10 | DMB: | Config problems for DMB: ME+2/1/10, CFEB4 5.0V reads 4.44
585 |  |  |  | 2008-07-22 |  |  | CC | VMEp2_3 |  | ME+2/1/10, | DMB: | Config problems for
480 |  |  |  | 2008-07-09 |  |  | CC | VMEp2_3 |  | ME+2/1/10, |  | Config problems for DMB:
411 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp2_5 |  | ME+2/1/15 |  | hot comparator channel CFEB4 layer 2 (old, minor)
140 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp2_5 |  | ME+2/1/15 |  | ME+2/1/15 – hot comparator channel CFEB4 layer 2 (old, minor)
586 |  |  |  | 2008-07-22 |  |  | CC | VMEp2_1 |  | ME+2/1/3, | ALCT: | Config problems for
487 |  |  |  | 2008-07-09 |  |  | Alex |  |  | ME+2/2/06 |  | Master #19 replaced (ME+2/2/06) |  |  |  |  | XXX
642 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+2/2/14 |  | remeasured timings, changed alct_rx = 11 -> 9, alct_tx = 8 -> 6
526 | 2 | study |  | 2008-07-18 |  |  | Misha | VMEp2_2 |  | ME+2/2/14 | ALCT | Sometimes ALCT bad, replace ALCT | 90 | 2h | access | Y
370 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEp2_2 |  | ME+2/2/14 | RUI#01, DDU Input 2: | ALCT NotPresent Error set 63 times
383 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEp2_2 |  | ME+2/2/14 | RUI#01, DDU Input 2: | ALCT NotPresent Error set 512 times
8 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_2 |  | ME+2/2/14 | ALCT | (data format problem)
216 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_2 |  | ME+2/2/14 | ALCT | data format problem: sometimes ALCT Not Present | 20% | 2h | cherry picker only. Next replace TMB/RAT
545 | 3 |  |  | 2008-07-20 |  |  | Misha | VMEp2_3 |  | ME+2/2/15 | CFEB5 | 80MHz mux clock dead, the SCA data path OK |  |  |  | Y
610 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+2/2/15 | CFEB#5 | 80MHz mux clock dead, the SCA data path OK |  |  |  |  | XXX
409 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp2_3 |  | ME+2/2/15 |  | no comparators data for CFEB5 (old)
138 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp2_3 |  | ME+2/2/15 |  | ME+2/2/15 – no comparators data for CFEB5 (old)
89 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEp2_3 |  | ME+2/2/15 | CFEB | - one “dead” CFEB?
325 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_3 |  | ME+2/2/15 | CFEB5 | (dead comps.) | 0.2 | 5h | remove chamber. Next replace TMB/DMB
569 |  |  |  | 2008-07-22 |  |  | jason | VMEp2_3 |  | ME+2/2/15, |  | replace CFEB5
527 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEp2_6 |  | ME+2/2/34 |  | HV Off for whole chamber; Now Fixed |  |  |  | Y
486 |  |  |  | 2008-07-09 |  |  | Alex | VMEp2_6 |  | ME+2/2/34 |  | Master #20 replaced (ME+2/2/34)
367 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEp2_6 |  | ME+2/2/34 |  | Empty CSC list
336 |  |  |  | 2008-07-06 |  |  | Madorsky | VMEp2_6 |  | ME+2/2/34 |  | MB020 was responsible for that. Recalibrated, should not happen anymore.
339 |  |  |  | 2008-07-06 |  |  | Misha | VMEp2_6 |  | ME+2/2/34 |  | HV Off for whole chamber;
190 |  |  |  | 2008-07-02 |  |  | Misha | VMEp2_6 |  | ME+2/2/34 |  | dead - HV? |  |  | review after next run
6 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_6 |  | ME+2/2/34 | ALCT | (Full FIFO @DMB) |  | 2h | cherry picker only
7 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_6 |  | ME+2/2/34 | CFEB5 | (Hot CFEB) |  | "free" | ^^ fix with above^^
214 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_6 |  | ME+2/2/34 | ALCT | (Full FIFO @DMB) |  | 2h | cherry picker only. Swapped TMB/RAT yet? Only 4 runs show this error, 47009-47062 during Cruzet2, looks OK now, why?
215 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_6 |  | ME+2/2/34 | CFEB5 | (Hot CFEB) | 50% | "free" | ^^ fix with above?^^
323 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_6 |  | ME+2/2/34 | ALCT | (Full FIFO @DMB) |  | 2h | cherry picker only. Swapped TMB/RAT yet? Only 4 runs show this error, 47009-47062 during Cruzet2, looks OK now, why?
324 |  |  |  | 2008-06-18 |  |  | Jason | VMEp2_6 |  | ME+2/2/34 | CFEB5 | (Hot CFEB) | 0.5 | "free" | ^^ fix with above?^^
570 |  |  |  | 2008-07-22 |  |  | jason | VMEp2_6 |  | ME+2/2/34, |  | replace CFEB5
629 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+3/1/04 |  | 0 cfeb5--kill comparator inputs (enableCLCTInputs_reg42=15)
528 | 2 | access |  | 2008-07-18 |  |  | Misha | VMEp3_1 |  | ME+3/1/04 | CFEB | Hot comparator - check cables at chamber - exclude CFEB until fixed | 90 | 2h | access to chamber | Y
462 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp3_1 |  | ME+3/1/04 |  | Hot comparator
402 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp3_1 |  | ME+3/1/04 |  | HOT, comparators CFEB5, layers 2,3 (TMB cable?) , 14% B-Words, ~7% CFEB DAV and CFEB CRC errors, if we can not fix it, mask CFEB5 [high]
62 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEp3_1 |  | ME+3/1/04 | CFEB5 | – Hot comparators CFEB5, layers 2,3 (TMB cable?)
10 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_1 |  | ME+3/1/04 | CFEB5 | (Hot CFEB) |  | 2h | cherry picker only
218 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_1 |  | ME+3/1/04 | CFEB5 | (Hot CFEB) | 20% | 2h | cherry picker only. Replaced TMB 21.04.08 10:00. replace DMB next
134 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp3_1 |  | ME+3/1/04 |  | HOT, comparators CFEB5, layers 2,3 (TMB cable?) , 14 CFEB DAV and CFEB CRC errors, if we can not fix it, mask CFEB5.
90 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEp3_1 |  | ME+3/1/04 - |  | “bad”; wires, strips – UP; rechits – DOWN
406 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp3_3 |  | ME+3/1/09 |  | no comparators data for CFEB1 (old)
135 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp3_3 |  | ME+3/1/09 |  | ME+3/1/09 – no comparators data for CFEB1 (old)
550 | 4 |  |  | 2008-07-20 |  |  | Misha | VMEp3_3 |  | ME+3/1/10 | ALCT | hot (trigger pass always "ON") |  |  |  | Y
322 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_3 |  | ME+3/1/10 | ALCT | (config problem) | 0.2 | 2h | cherry picker only. Next replace TMB/RAT
568 |  |  |  | 2008-07-22 |  |  | jason | VMEp3_3 |  | ME+3/1/10, |  | replace ALCT (from Misha)
407 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp3_4 |  | ME+3/1/11 |  | noisy SCA channel CFEB1 layer 1 (old, minor)
136 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp3_4 |  | ME+3/1/11
625 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+3/1/17 |  | kill at DDU setting kill_input=3 to kill cfeb1 was not sufficient. (timing OK)
587 |  |  |  | 2008-07-22 |  |  | CC | VMEp3_6 |  | ME+3/1/17 | DMB: | Config problems for
530 | 4 | study |  | 2008-07-21 |  |  | Misha | VMEp3_6 |  | ME+3/1/17 | ALCT | data format problem: ALCT Not Present - present 18-July run | 0.2 | 2h | remove 2+ CSCs. | Y
552 |  |  |  | 2008-07-21 |  |  | Greg | VMEp3_6 |  | ME+3/1/17 | CFEB1 | kill_input=3
608 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+3/1/17 | CFEB#1 | FPGA not programmed |  |  |  |  | XXX
529 | 3 | remove |  | 2008-07-18 |  |  | Jason | VMEp3_6 |  | ME+3/1/17 | CFEB1 | dead comps. & SCA, configure problem - under investigation - ALCT? | 20% | 8h | remove 2+ CSCs. | Y
514 |  |  |  | 2008-07-17 |  |  | CC | VMEp3_6 |  | ME+3/1/17 |  | Config problems for DMB ME+3/2/35
463 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEp3_6 |  | ME+3/1/17 |  | Bad ALCT and CFEB1
354 |  |  |  | 2008-07-06 |  |  | CC | VMEp3_6 |  | ME+3/1/17 | DMB: | Config problems for
319 |  |  |  | 2008-07-04 | 0.48958 |  | check config | VMEp3_6 |  | ME+3/1/17 | DMB: | Config problems for
304 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEp3_6 |  | ME+3/1/17 | DMB
268 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp3_6 |  | ME+3/1/17 | DMB | Config problems for DMB
408 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp3_6 |  | ME+3/1/17 |  | (was masked) 40-70% CFEB, TMB, ALCT DAV errors and DMB Trailer missing errors, no SCA data for CFEB1
115 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEp3_6 |  | ME+3/1/17 | DMB | Config problems for DMB: ME+3/1/17,
4 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_6 |  | ME+3/1/17 | CFEB1 | (dead comps. & SCA) |  | 8h | remove 2+ CSCs
212 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_6 |  | ME+3/1/17 | ALCT | data format problem: ALCT Not Present | 20% | 2h | cherry picker only. Next replace TMB/RAT
213 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_6 |  | ME+3/1/17 | CFEB1 | dead comps. & SCA, configure problem | 20% | 8h | remove 2+ CSCs. Next replace TMB/DMB
321 |  |  |  | 2008-06-18 |  |  | Jason | VMEp3_6 |  | ME+3/1/17 | ALCT | data format problem: ALCT Not Present | 0.2 | 2h | cherry picker only. Next replace TMB/RAT
51 | [6 empty] | Check Config | VMEp3_6 |  | ME+3/1/17 | DMB: | Config problems for DMB: ME+3/1/17, CFEB1 5.0V I=0.73 (instead of something more like 0.92)... (<- but this is not that big of a deal says Misha). Went to DMB Utils DMB Test All. CFEB5 reads wrong comparator (<- this apparently is the real problem).
137 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp3_6 |  | ME+3/1/17 |  | (was masked) 40-70
481 |  |  |  | 2008-07-09 |  |  | CC | VMEp3_6 |  | ME+3/1/17, |  | Config problems for DMB:
405 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEp3_3 |  | ME+3/2/15 |  | few noisy comparators and SCA channels CFEB5 layer 6 (old, flaky, minor)
133 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEp3_3 |  | ME+3/2/15 |  | ME+3/2/15 – few noisy comparators and SCA channels CFEB5 layer 6 (old, flaky, minor)
555 | 2 | study |  | 2008-07-20 |  |  | Misha | VMEp3_4 |  | ME+3/2/21 | CFEB3 | bad CRCs |  |  |  | Y
640 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+3/2/24 |  | remeasured timings, changed tmb_lct_cable_delay from 3 -> 2
630 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME+3/2/35 |  | 6 cfeb4
554 | 3 |  |  | 2008-07-21 |  |  | Greg | VMEp3_6 |  | ME+3/2/35 | CFEB4 | kill_input=6 |  |  |  | Y
611 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME+3/2/35 | CFEB#4 | FPGA not programmed |  |  |  |  | XXX
573 |  |  |  | 2008-07-22 |  |  | jason | VMEp3_6 |  | ME+3/2/35, |  | replace CFEB4 (Gu)
588 |  |  |  | 2008-07-22 |  |  | CC | VMEp3_6 |  | ME+3/2/35, | DMB: | Config problems for
269 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEp4_5 |  | ME+4/1/15 | DMB | Config problems for ALCT
91 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEp4_3 |  | ME+4/1/8-10 |  | “bad”? feature?
428 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_1 |  | ME-1/1/01 |  | ~0.5% CFEB DAV, Sample count errors [low]
152 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_1 |  | ME-1/1/01 |  | ~0.5
531 | 2 | remove |  | 2008-07-18 |  |  | Jason | VMEm1_2 |  | ME-1/1/03 | CFEB5 | hot comparator pl 1,4 - one channel hot only | 80 | 8h | pull chamber | Y
193 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_2 |  | ME-1/1/03 | CFEB5 | hot comparator pl 1,4 |  | 8h | pull chamber
429 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_2 |  | ME-1/1/03 |  | hot comparators channel CFEB5 layers 1,4 [low]
153 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_2 |  | ME-1/1/03 |  | ME-1/1/03 – hot comparators channel CFEB5 layers 1,6
631 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-1/1/04 |  | 1 ALCT -> fuse?
545 |  |  |  | 2008-07-21 |  |  | Greg | VMEm1_2 |  | ME-1/1/04 | ALCT | kill_input=1
532 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEm1_2 |  | ME-1/1/04 | ALCT | Fuse out | 100 | 8h | pull chamber | Y
492 |  |  |  | 2008-07-17 |  |  | CC | VMEm1_2 |  | ME-1/1/04 |  | ALCT firmware incorrect
464 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEm1_2 |  | ME-1/1/04 |  | Bad ALCT
355 |  |  |  | 2008-07-06 |  |  | CC | VMEm1_2 |  | ME-1/1/04 | ALCT: | Config problems for
194 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_2 |  | ME-1/1/04 | ALCT | Fuse out | 100 | 8h | pull chamber
285 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEm1_2 |  | ME-1/1/04 | ALCT
286 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEm1_2 |  | ME-1/1/04 | ALCT
252 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm1_2 |  | ME-1/1/04 | ALCT | Config problems for ALCT
272 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm1_2 |  | ME-1/1/04 |  | ALCT firmware incorrect
442 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_2 |  | ME-1/1/04 |  | Very Hot, 100% ALCT DAV errors for 280K events and DMB-Input timeout for ALCT data [high]
116 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm1_2 |  | ME-1/1/04 | ALCT | Config problems for ALCT: ME-1/1/4,
52 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm1_2 |  | ME-1/1/04 | ALCT: | Config problems for ALCT: ME-1/1/4, Check TMB+ALCT fixed problem
66 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_2 |  | ME-1/1/04 | ALCT | – Very Hot, ALCT DAV errors for 280K events and DMB-Input timeout for ALCT data
92 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm1_2 |  | ME-1/1/04 |  | “dead”; no wires, no strips
15 |  |  |  | 2008-06-18 |  |  | Misha | VMEm1_2 |  | ME-1/1/04 | ALCT | TMB slot 4, ALCT 1.8 V fuse blown |  | 5h | remove chamber
154 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_2 |  | ME-1/1/04 |  | Very Hot, 100
598 |  |  |  | 2008-07-22 |  |  | CC | VMEm1_2 |  | ME-1/1/04, | ALCT: | Config problems for
469 |  |  |  | 2008-07-09 |  |  | CC | VMEm1_2 |  | ME-1/1/04, |  | Config problems for ALCT:
243 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_2 |  | ME-1/1/05 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors, had Config problems. |  | done | not seen after 21 June 08, Config problems were resolved
417 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_2 |  | ME-1/1/05 |  | 80% events all 5 CFEBs report DMB Input FIFO full + B-words
67 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_2 |  | ME-1/1/05 | CFEB | - 80% events all 5 CFEBs report DMB Input FIFO full + B-words
93 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm1_2 |  | ME-1/1/05 | CFEB | “dead”; CFEB problems? (wires look fine)
39 |  |  |  | 2008-06-18 |  |  | Jason | VMEm1_2 |  | ME-1/1/05 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors, had Config problems. GONE |  | done | not seen after 21 June 08, Config problems were resolved
155 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_2 |  | ME-1/1/05 |  | 80
305 |  |  |  | 2008-07-03 |  |  | Misha | VMEm1_3 |  | ME-1/1/06 | CFEB | Fixed
195 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_3 |  | ME-1/1/06 | CFEB1 & 3 | dead, comparator OK | 100 | 8h | pull chamber
244 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_3 |  | ME-1/1/06 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors |  |  | replaced DMB 06.30.08 16:15, still bad. CHECK TIMING
430 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_3 |  | ME-1/1/06 |  | (was missing in last reported run) 60-85% events all 5 CFEBs report DMB Input FIFO full + B-words [medium]
156 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_3 |  | ME-1/1/06 |  | (was missing in last reported run) 60-85
632 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-1/1/15 |  | 1 checked timing to be OK (CFEB 4 rx not very clean--maybe need TMB swap?)
356 |  |  |  | 2008-07-06 |  |  | CC | VMEm1_7 |  | ME-1/1/18, ME-1/1/19, ME-1/1/20, ME-1/2/18, ME-1/2/19, ME-1/2/20, ME-1/3/18, ME-1/3/19, ME-1/3/20, | ALCT: | Config problems for
196 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_7 |  | ME-1/1/19 |  | comparator strange
431 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_7 |  | ME-1/1/19 |  | 1% ALCT DAV errors, 0.5% CFEB CRC and Sample count errors [low]
157 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_7 |  | ME-1/1/19 |  | 1 CFEB CRC and Sample count errors
395 |  |  |  | 2008-07-08 |  |  | Vladimir Tchkhovski | VMEm1_8 |  | ME-1/1/22 | HV | /Layer5: VOSET = 500V (Status = "OFF")
396 |  |  |  | 2008-07-08 |  |  | Vladimir Tchkhovski | VMEm1_8 |  | ME-1/1/22 | HV | /Layer6: VOSET = 500V (Status = "OFF")
483 |  |  |  | 2008-07-08 |  |  | Tchekhovski | VMEm1_8 |  | ME-1/1/22 | HV | /Layer5: VOSET = 500V (Status = "OFF")
484 |  |  |  | 2008-07-08 |  |  | Tchekhovski | VMEm1_8 |  | ME-1/1/22 | HV | /Layer6: VOSET = 3000V (Status = "ON")
197 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_8 |  | ME-1/1/22 | ALCT | setup? CFEB5
245 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_8 |  | ME-1/1/22 | CFEB5 | CFEB5 hot comparators |  |  | next: do timing scan, check TMB cable, replace TMB or CFEB5
443 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_8 |  | ME-1/1/22 |  | hot comparators CFEB5 (TMB cable?), low anode occupancy in group of channels in all layers (see “other issues 2” plot) [medium]
158 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_8 |  | ME-1/1/22 |  | hot comparators CFEB5 (TMB cable?), low anode occupancy in group of channels in all layers (thresholds?) (see “other issues 2” plot)
533 | 3 | remove |  | 2008-07-18 |  |  | Misha | VMEm1_9 |  | ME-1/1/25 | HV | HV channel will not stay at full voltage - 2800 | 90 | 8h | disable HV channel - replace chamber | Y
397 |  |  |  | 2008-07-07 |  |  | Vladimir Tchkhovski | VMEm1_9 |  | ME-1/1/25 | HV | /layer4: V0SET = 2800V ("nominal" - 200V)
307 |  |  |  | 2008-07-03 |  |  | Vladimir | VMEm1_9 |  | ME-1/1/25 | HV | with the exception of ME-1/1/25/Layer5 ("I-tripping")
432 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_9 |  | ME-1/1/25 |  | 11% CFEB CRC errors, 6% CFEB DAV and Sample count errors [medium]
159 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_9 |  | ME-1/1/25 |  | 11 CFEB DAV and Sample count errors
250 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_9 |  | ME-1/1/26 | ALCT | ALCT problems: few CRC errors, not seen by TMB |  |  | must be problem on TMB, DMB or backplane
334 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_9 |  | ME-1/1/26 | ALCT | ALCT problems: few CRC errors, not seen by TMB |  |  | must be problem on TMB, DMB or backplane
433 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_9 |  | ME-1/1/26 |  | <0.5% ALCT CRC errors (starting from run 080624_090805) [low]
160 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_9 |  | ME-1/1/26 |  | <0.5
543 | 1 | replace |  | 2008-07-20 |  |  | Jianhui | VMEm1_10 |  | ME-1/1/29 | CFEB5 | JTAG TDO read problem |  |  |  | Y
605 |  |  |  | 2008-07-20 |  |  | Misha | VMEm1_10 |  | ME-1/1/29 | CFEB#5 | JTAG TDO read problem
489 |  |  |  | 2008-07-15 |  |  | Jianhui | VMEm1_10 |  | ME-1/1/29 |  | CFEB#5 on ME-1/1/29 (VME crate VME-1/10, alot #7) has problem with JTAG TDO. The CFEB seems to be working fine, but the JTAG readback is not correct. By swapping the DMB inputs, the problem is confirmed on the CFEB#5 or its cable. There is no need to disable that CFEB (or chamber), it only causes 'check configure' problem for now.
357 |  |  |  | 2008-07-06 |  |  | CC | VMEm1_10 |  | ME-1/1/29 | DMB: | Config problems for
418 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_10 |  | ME-1/1/29 |  | Gap in SCA and comparators data CFEB5
161 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_10 |  | ME-1/1/29 |  | ME-1/1/29 – Gap in SCA and comparators data CFEB7
53 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm1_10 |  | ME-1/1/29 | DMB: | Config problems for DMB ME-1/1/29 lost communication with CFEB5 (this is not new). Config problems for DMB for ME-1/2/28... went to DMB utils, clicked DMB CONTROL Load Firmware, Write Flash to Crate, Power-up Init.
574 |  |  |  | 2008-07-22 |  |  | jason | VMEm1_10 |  | ME-1/1/29, |  | replace CFEB5 (Gu, lower priority)
589 |  |  |  | 2008-07-22 |  |  | CC | VMEm1_10 |  | ME-1/1/29, | DMB: | Config problems for
491 |  |  |  | 2008-07-17 |  |  | CC | VMEm1_11 |  | ME-1/1/32 |  | Config problems for DMB
590 |  |  |  | 2008-07-22 |  |  | CC | VMEm1_11 |  | ME-1/1/32, | DMB: | Config problems for
468 |  |  |  | 2008-07-09 |  |  | CC | VMEm1_11 |  | ME-1/1/32, |  | Config problems for DMB:
246 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_12 |  | ME-1/1/33 | CFEB5 | minor CFEB5 problem: often bad CRCs |  |  | next: do timing scan, replace CFEB5
419 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_12 |  | ME-1/1/33 |  | 0.5% CFEB CRC (DMB was replaced)
68 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_12 |  | ME-1/1/33 | CFEB | - 0.5% CFEB CRC (minor)
40 |  |  |  | 2008-06-18 |  |  | Jason | VMEm1_12 |  | ME-1/1/33 | CFEB5 | minor CFEB5 problem: often bad CRCs |  |  | next: check timing scan, replace CFEB5
162 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_12 |  | ME-1/1/33 |  | 0.5
359 |  |  |  | 2008-07-06 |  |  | CC | VMEm1_12 |  | ME-1/1/33, ME-1/1/34, ME-1/1/35, ME-1/2/33, ME-1/2/34, ME-1/2/35, ME-1/3/33, ME-1/3/34, ME-1/3/35, | ALCT: | Config problems for
465 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEm1_12 |  | ME-1/1/35 |  | Bad CFEB4
534 | 3 | study |  | 2008-07-09 |  |  | Jason | VMEm1_12 |  | ME-1/1/35 | CFEB4 | Bad CFEB4 needs more study | 20 | 5h |  | Y
488 |  |  |  | 2008-07-15 |  |  | Jianhui | VMEm1_1 |  | ME-1/1/36 |  | The ME-1/1/36 dead channels are caused by a software bug. Right now, the software bug is fixed. I am committing the code to CVS. There is NO dead channel on ME-1/1/36 calibration runs.
198 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_1 |  | ME-1/1/36 | DMB | structure |  |  | DMB swap
600 |  |  |  | 2008-07-20 |  |  | Misha | VMEm1_2 |  | ME-1/10/4 | ALCT i | s dead
493 |  |  |  | 2008-07-17 |  |  | CC | VMEm1_2 |  | ME-1/10/4 |  | Config problems for ALCT
560 |  |  |  | 2008-07-22 |  |  | Luca | VMEm1_2 |  | ME-1/2/03 | ALCT | reading bad firmware & 1.8 V I~0.1 A. Power cycling and hard reset didn't help
35 |  |  |  | 2008-06-18 |  |  | Jason | VMEm1_3 |  | ME-1/2/08 | CFEB3 | CFEB3 problem: DAV-LCT error one time, GONE |  | done | cable check resolved problem
240 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_6 |  | ME-1/2/17 | CFEB5 | CFEB5 problem: one run set Not Present |  | done | not seen after 18 June 08 when cable was checked
331 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_6 |  | ME-1/2/17 | CFEB5 | CFEB5 problem: one run set Not Present |  | done | not seen after 18 June 08 when cable was checked
104 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm1_6 |  | ME-1/2/17 |  | wires “hot”; (ME-1/2/17 is “masked”)
36 |  |  |  | 2008-06-18 |  |  | Jason | VMEm1_6 |  | ME-1/2/17 | CFEB5 | CFEB5 problem: one run set Not Present, GONE |  | done | not seen after 18 June 08 when cable was checked
248 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_7 |  | ME-1/2/19 | ALCT | ALCT problems: Not Present and L1A mismatch; often since 24.06.08 |  |  | check ALCT cables, do phase/timing scan
434 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_7 |  | ME-1/2/19 |  | ALCT L1A out of sync (messed up) [medium]
163 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_7 |  | ME-1/2/19 |  | ALCT L1A out of sync (messed up)
435 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_8 |  | ME-1/2/22 |  | no TMB data, noisy SCA channels CFEB3 layer 3,5 [high]
69 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_8 |  | ME-1/2/22 | CFEB5 | – Hot comparators CFEB5 (TMB cable?)
164 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_8 |  | ME-1/2/22 |  | no TMB data, noisy SCA channels CFEB3 layer 3,5
241 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_9 |  | ME-1/2/24 | CFEBs | all CFEBs have DAV-LCT errors |  | done | not seen after 24 June 08, timing tuned 30 June ~3pm, fixed it.
420 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_9 |  | ME-1/2/24 |  | all 5 CFEBs report no SCA data
70 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_9 |  | ME-1/2/24 | CFEB | - all 5 CFEBs report no SCA data
94 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm1_9 |  | ME-1/2/24 | CFEB | - “dead”; CFEB problems? (wires look fine)
37 |  |  |  | 2008-06-18 |  |  | Jason | VMEm1_9 |  | ME-1/2/24 | CFEBs | all CFEBs have DAV-LCT errors |  |  | not seen after 24 June 08, timing tuned 30 June ~3pm
165 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_9 |  | ME-1/2/24 |  | all 5 CFEBs report no SCA data
242 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_10 |  | ME-1/2/28 | CFEB2 | CFEB2 problem: Not Present and CRCs, etc. |  | done | not seen after 21 June 08, replaced DMB 23 June 9:30am
421 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_10 |  | ME-1/2/28 |  | 15% CFEB DAV Errors and 10% DMB-Input timeout for CFEB2
71 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_10 |  | ME-1/2/28 | CFEB | - 15% CFEB DAV Errors and 10% DMB-Input timeout for CFEB2
38 |  |  |  | 2008-06-18 |  |  | Jason | VMEm1_10 |  | ME-1/2/28 | CFEB2 | CFEB2 problem: Not Present and CRCs, etc. GONE |  | done | not seen after 21 June 08, replaced DMB 23 June 9:30am
166 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_10 |  | ME-1/2/28 |  | 15 DMB-Input timeout for CFEB2
490 |  |  |  | 2008-07-17 |  |  | CC | VMEm1_11 |  | ME-1/2/30 |  | DMB VME firmware incorrect
282 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEm1_11 |  | ME-1/2/30 | DMB VME
270 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm1_11 |  | ME-1/2/30 |  | DMB VME firmware incorrect
437 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_12 |  | ME-1/2/35 |  | no ALCT and TMB data [high]
170 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_12 |  | ME-1/2/35 |  | ME-1/2/35 – no ALCT and TMB data
199 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_2 |  | ME-1/3/03 | CFEB1 & 3 | hot comparator
444 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_2 |  | ME-1/3/03 |  | Hot comparator channels layers 1,2,3,5,6 border CFEB1-CFEB2 [low]
72 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm1_2 |  | ME-1/3/03 | CFEB | – Hot comparator channels layers 1,2,3,5,6 border CFEB1-CFEB2
167 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_2 |  | ME-1/3/03 |  | ME-1/3/03 – Hot comparator channels layers 1,2,3,5,6 border CFEB1-CFEB4
535 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEm1_6 |  | ME-1/3/17 | HV | HV Off for whole chamber; Now Fixed |  |  |  | Y
337 |  |  |  | 2008-07-06 |  |  | Madorsky | VMEm1_6 |  | ME-1/3/17 |  | MB026 was responsible for both of these. Also recalibrated.
340 |  |  |  | 2008-07-06 |  |  | Misha | VMEm1_6 |  | ME-1/3/17 |  | HV Off for whole chamber;
191 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_6 |  | ME-1/3/17 |  | dead - HV |  |  | HV action needed
536 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEm1_7 |  | ME-1/3/18 | HV | HV Off for whole chamber; Now Fixed |  |  |  | Y
338 |  |  |  | 2008-07-06 |  |  | Madorsky | VMEm1_7 |  | ME-1/3/18 |  | MB026 was responsible for both of these. Also recalibrated.
341 |  |  |  | 2008-07-06 |  |  | Misha | VMEm1_7 |  | ME-1/3/18 |  | HV Off for whole chamber;
192 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_7 |  | ME-1/3/18 |  | dead - HV? |  |  | review after next run
561 |  |  |  | 2008-07-22 |  |  | Luca | VMEm1_11 |  | ME-1/3/31 | ALCT | No communication. Readings wrong. TMB was replaced (5015 in 5260 out) but did not help. RAT boards were swapped.
537 | 4 | study |  | 2008-07-18 |  |  | Misha | VMEm1_11 |  | ME-1/3/31 | ALCT | ALCT was swapped, but problem remains, problem unknown | 100 | 2h | ALCT swap | Y
358 |  |  |  | 2008-07-06 |  |  | CC | VMEm1_11 |  | ME-1/3/31 | ALCT: | Config problems for
200 |  |  |  | 2008-07-02 |  |  | Misha | VMEm1_11 |  | ME-1/3/31 | ALCT | Dead | 100 | 2h | ALCT swap
249 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_11 |  | ME-1/3/31 | ALCT | ALCT problems: Full@DMB, Not Present, Timeout; often since 30.06.08. configure problem |  |  | check ALCT cables, do phase/timing scan
283 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEm1_11 |  | ME-1/3/31 | ALCT
284 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEm1_11 |  | ME-1/3/31 | ALCT
333 |  |  |  | 2008-07-02 |  |  | Jason | VMEm1_11 |  | ME-1/3/31 | ALCT | ALCT problems: Full@DMB, Not Present, Timeout; often since 30.06.08. configure problem |  |  | check ALCT cables, do phase/timing scan
251 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm1_11 |  | ME-1/3/31 | ALCT | Config problems for ALCT
271 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm1_11 |  | ME-1/3/31 |  | ALCT firmware incorrect
436 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_11 |  | ME-1/3/31 |  | 100% ALCT DMB Input FIFO full (timing?) [medium]
117 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm1_11 |  | ME-1/3/31 | ALCT | Config problems for ALCT: ME-1/3/31,
54 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm1_11 |  | ME-1/3/31 | ALCT: | Config problems for ALCT: ME-1/3/31... no communication with slow control. Not big problem, but not good in any case. Tried to power cycle CRB on this crate and then did Power-up Init, but still no good. Tried CCB hard reset, no good. Went to DM status, A
16 |  |  |  | 2008-06-18 |  |  | Misha | VMEm1_11 |  | ME-1/3/31 | ALCT | ALCT on ME-1/3/31 was replaced but still has a problem of occasional reading of slow control parameters, sometimes no ALCT triggers.
168 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_11 |  | ME-1/3/31 |  | 100
591 |  |  |  | 2008-07-22 |  |  | CC | VMEm1_11 |  | ME-1/3/31, | ALCT: | Config problems for
422 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm1_11 |  | ME-1/3/32 |  | Strange SCA occupancy (saw that before, DMB cables swap?)
169 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm1_11 |  | ME-1/3/32 |  | ME-1/3/32 – Strange SCA occupancy (saw that before, DMB cables swap?)
385 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm2_2 |  | ME-2/1/06 | RUI#20, DDU Input 13: | ALCT L1A Error set 175 times
234 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_2 |  | ME-2/1/06 | ALCT | ALCT L1A Errors, 702 in one run |  | done | run 080621_1353: ALCT off by +1 from beginning to end
29 |  |  |  | 2008-06-18 |  |  | Jason | VMEm2_2 |  | ME-2/1/06 | ALCT | ALCT L1A Errors, 702 in one run, GONE |  | done | run 080621_1353: ALCT off by +1 from beginning to end
496 |  |  |  | 2008-07-17 |  |  | CC | VMEm2_3 |  | ME-2/1/09 |  | Config problems for ALCT
361 |  |  |  | 2008-07-06 |  |  | CC | VMEm2_3 |  | ME-2/1/09 | ALCT: | Config problems for
287 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEm2_3 |  | ME-2/1/09 | ALCT
254 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm2_3 |  | ME-2/1/09 | ALCT | Config problems for ALCT
126 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_3 |  | ME-2/1/09 | ALCT | Config problems for ALCT: ME-2/1/9,
55 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm2_3 |  | ME-2/1/09 | ALCT: | Config problems for ALCT: ME-2/1/9, clicked on Configure TMB+ALCT, this is well known problem.
235 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_3 |  | ME-2/1/10 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors |  | done | timing tuned 30 June 08, ~3pm, fixed it.
329 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_3 |  | ME-2/1/10 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors |  | done | timing tuned 30 June 08, ~3pm, fixed it.
423 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_3 |  | ME-2/1/10 |  | 90% events all 5 CFEBs report DMB Input FIFO full + B-words
73 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm2_3 |  | ME-2/1/10 | CFEB | - 90% events all 5 CFEBs report DMB Input FIFO full + B-words
95 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm2_3 |  | ME-2/1/10 | CFEB | CFEB problems? (wires look fine)
97 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm2_3 |  | ME-2/1/10 | CFEB | - “bad”; CFEB problems? (wires look fine)
30 |  |  |  | 2008-06-18 |  |  | Jason | VMEm2_3 |  | ME-2/1/10 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors, GONE? |  |  | timing tuned 30 June 08, ~3pm
171 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_3 |  | ME-2/1/10 |  | 90
438 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_4 |  | ME-2/1/11 |  | low occupancy and no anode data in many segments (see “other issues 4” plot) [medium]
118 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/11 | ALCT | Config problems for ALCT:
119 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/11 | DMB | Config problems for DMB:
172 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_4 |  | ME-2/1/11 |  | ME-2/1/11 – low occupancy and no anode data in many segments (see “other issues 4” plot)
120 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/12 | ALCT | Config problems for ALCT:
121 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/12 | TMB | Config problems for TMB:
122 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/12 | DMB | Config problems for DMB:
274 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm2_4 |  | ME-2/1/13 |  | DMB Control FPGA firmware incorrect
123 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/13 | ALCT | Config problems for ALCT:
124 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/13 | TMB | Config problems for TMB:
125 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/1/13 | DMB | Config problems for DMB:
273 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm2_4 |  | ME-2/1/13 CFEB 4,ME-2/1/13 CFEB 3,ME-2/1/13 CFEB 2,ME-2/1/13 CFEB 1,ME-2/1/13 CFEB 0,
592 |  |  |  | 2008-07-22 |  |  | CC | VMEm2_3 |  | ME-2/1/9, | ALCT: | Config problems for
470 |  |  |  | 2008-07-09 |  |  | CC | VMEm2_3 |  | ME-2/1/9, |  | Config problems for ALCT:
59 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm2_6 |  | ME-2/2/01 | DMB: | Config problems for DMB: ME-2/2/1, DMB fpga user ID reads all f’s. DMB CONTROL Load Firmware. Stan Durkin says
236 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_6 |  | ME-2/2/02 | CFEBs | bad CFEB efficiency |  | done | timing tuned 30 June 08, ~3pm, fixed it.
424 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_6 |  | ME-2/2/02 |  | 95% events all 5 CFEBs report DMB Input FIFO full + B-words
74 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm2_6 |  | ME-2/2/02 | CFEB | - 95% events all 5 CFEBs report DMB Input FIFO full + B-words
31 |  |  |  | 2008-06-18 |  |  | Misha | VMEm2_6 |  | ME-2/2/02 | CFEBs | bad CFEB efficiency, GONE? |  |  | timing tuned 30 June 08, ~3pm
173 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_6 |  | ME-2/2/02 |  | 95
96 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm2_6 |  | ME-2/2/02 | CFEB | CFEB problems? (wires look fine)
633 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-2/2/03 |  | 4 cfeb2
547 | 3 |  |  | 2008-07-21 |  |  | Greg | VMEm2_1 |  | ME-2/2/03 | CFEB2 | kill_input=4
603 |  |  |  | 2008-07-20 |  |  | Misha | VMEm2_1 |  | ME-2/2/03 | CFEB#2 | FPGA not programmed
538 | 3 | remove 3 |  | 2008-07-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/03 | CFEB2 | dead - must swap CFEB | 100 | 8h | swap >access, remove front chambers | Y
494 |  |  |  | 2008-07-17 |  |  | CC | VMEm2_1 |  | ME-2/2/03 | CFEB 1 | CFEB firmware incorrect
495 |  |  |  | 2008-07-17 |  |  | CC | VMEm2_1 |  | ME-2/2/03 |  | Config problems for DMB
485 |  |  |  | 2008-07-08 |  |  | Killewald | VMEm2_1 |  | ME-2/2/03 |  | Masked OFF for problems in DDU
381 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm2_1 |  | ME-2/2/03 | RUI#36, DDU Input 10: | CFEBs flagged with Bad DAV-LCT: 0x02 = CFEB #2
382 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm2_1 |  | ME-2/2/03 | RUI#36, DDU Input 10: | CFEB DAV-LCT match Error set 9 times
393 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm2_1 |  | ME-2/2/03 | RUI#36, DDU Input 10: | CFEBs flagged with Bad DAV-LCT: 0x02 = CFEB #2
394 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm2_1 |  | ME-2/2/03 | RUI#36, DDU Input 10: | CFEB DAV-LCT match Error set 64 times
360 |  |  |  | 2008-07-06 |  |  | CC | VMEm2_1 |  | ME-2/2/03 | DMB: | Config problems for
201 |  |  |  | 2008-07-02 |  |  | Misha | VMEm2_1 |  | ME-2/2/03 | CFEB2 | dead | 100 | 8h | swap >access, remove front chambers
247 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_1 |  | ME-2/2/03 | CFEB2 | CFEB2 problem: DAV-LCT, Full@DMB errors; often since 20.06.08 |  |  | check DMB cable, do timing scan.
332 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_1 | VMEm2_1 | ME-2/2/03 | CFEB2 | CFEB2 problem: DAV-LCT, Full@DMB errors; often since 20.06.08 |  |  | check DMB cable, do timing scan.
253 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm2_1 | VMEm2_1 | ME-2/2/03 | DMB | Config problems for DMB
439 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_1 |  | ME-2/2/03 |  | 94% CFEB2 DMB Input FIFO full (timing?) [medium]
482 |  |  |  | 2008-07-09 |  |  | CC | VMEm2_1 |  | ME-2/2/03 |  | VMEm2_1, config problems for DMB:ME-2/2/3
174 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_1 |  | ME-2/2/03 |  | 94
593 |  |  |  | 2008-07-22 |  |  | CC | VMEm2_1 |  | ME-2/2/03, | DMB: | Config problems for
539 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/05 | HV | – low anode occupancy in layer 2 segment (minor, HV) - Fixed |  |  |  | Y
75 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm2_1 |  | ME-2/2/05 |  | – low anode occupancy in layer 2 segment (minor, HV)
540 | 2 | remove |  | 2008-07-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/06 | HV | one HV sector constant trip; Alex states the problem in in the chamber | 90 | 8h | replace chamber | Y
342 |  |  |  | 2008-07-06 |  |  | Misha | VMEm2_1 |  | ME-2/2/06 |  | one HV sector constant trip;
202 |  |  |  | 2008-07-02 |  |  | Misha | VMEm2_1 |  | ME-2/2/06 | HV sector | off |  |  | HV action needed
454 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_1 |  | ME-2/2/06 |  | tripped HV segment, no anode data (minor)
175 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_1 |  | ME-2/2/06 |  | ME-2/2/06 – tripped HV segment, no anode data (minor)
32 |  |  |  | 2008-06-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/07 | ALCT | no VME reading back status of TMB or ALCT |  |  | replaced TMB 06.19.08 16:20
238 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_1 |  | ME-2/2/08 | CFEB | bad occupancy for CFEBs, 10 times ALCT rate, even have CLCTs with no HV |  | done | replaced TMB 06.19.08 12:20, then timing tuned 30 June ~3pm
425 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_1 |  | ME-2/2/08 |  | very low CFEB SCA data efficiency for all 5 CFEBs
76 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm2_1 |  | ME-2/2/08 | CFEB | - very low CFEB SCA data efficiency for all 5 CFEBs
33 |  |  |  | 2008-06-18 |  |  | Misha | VMEm2_1 |  | ME-2/2/08 | CFEB | bad occupancy for CFEBs, 10 times ALCT rate, even have CLCTs with no HV |  |  | replaced TMB 06.19.08 12:20, then timing tuned 30 June ~3pm
176 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_1 |  | ME-2/2/08 |  | very low CFEB SCA data efficiency for all 5 CFEBs
372 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm2_3 |  | ME-2/2/20 | RUI#22, DDU Input 2: | ALCT L1A Error set 31 times
386 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm2_3 |  | ME-2/2/20 | RUI#22, DDU Input 2: | ALCT L1A Error set 229 times
239 |  |  |  | 2008-07-02 |  |  | Jason | VMEm2_3 |  | ME-2/2/20 | ALCT | ALCT L1A Errors, continually gets farther behind |  |  | next: do timing scan, replace TMB, RAT or ALCT
445 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_3 |  | ME-2/2/20 |  | ALCT L1A out of sync - really messed up [medium]
34 |  |  |  | 2008-06-18 |  |  | Jason | VMEm2_3 |  | ME-2/2/20 | ALCT | ALCT L1A Errors, continually gets farther behind |  |  | next: check timing scan, replace TMB, RAT or ALCT
177 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_3 |  | ME-2/2/20 |  | ALCT L1A out of sync
440 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm2_4 |  | ME-2/2/21 |  | 95% CFEB4 DMB Input FIFO full (timing?) [medium]
127 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_4 |  | ME-2/2/21 | DMB | Config problems for DMB:
56 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm2_4 |  | ME-2/2/21 | DMB: | Config problems for DMB: ME-2/2/21, CFEB4 lost communication with slow control. DMB turn off LV, then DMB turn on LV. CFEB4 reads low voltage.
178 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm2_4 |  | ME-2/2/21 |  | 95
562 |  |  |  | 2008-07-22 |  |  | Luca | VMEm2_4 |  | ME-2/2/22 | DMB | chip1 1.8 V reads as 1843 mV
77 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm2_4 |  | ME-2/2/22 |  | ALCT L1A out of sync - really messed up
634 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-2/2/24 |  | kill at DDU timing checked to be OK
57 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm2_5 |  | ME-2/2/27 | DMB: | Config problems for DMB: ME-2/2/27, CFEB1 lost communication with slow control
362 |  |  |  | 2008-07-06 |  |  | CC | VMEm2_6 |  | ME-2/2/35 | ALCT: | Config problems for
635 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-2/2/36 |  | 1 ALCT -> fuse
544 | 4 |  |  | 2008-07-21 |  |  | Greg | VMEm2_6 |  | ME-2/2/36 | ALCT | kill_input=1
599 |  |  |  | 2008-07-20 |  |  | Misha | VMEm2_6 |  | ME-2/2/36 | ALCT i | s dead
17 | 4 | access |  | 2008-07-18 |  |  | Misha | VMEm2_6 |  | ME-2/2/36 | ALCT | TMB slot 16 ALCT 1.8 V fuse blown. | 90 | 5h | need YE-2 open | Y
497 |  |  |  | 2008-07-17 |  |  | CC | VMEm2_6 |  | ME-2/2/36 |  | ALCT firmware incorrect
498 |  |  |  | 2008-07-17 |  |  | CC | VMEm2_6 |  | ME-2/2/36 |  | Config problems for ALCT
368 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm2_6 |  | ME-2/2/36 |  | Empty CSC list
363 |  |  |  | 2008-07-06 |  |  | CC | VMEm2_6 |  | ME-2/2/36 | ALCT: | Config problems for
288 |  |  |  | 2008-07-02 |  | 080702_141215 | check config | VMEm2_6 |  | ME-2/2/36 | ALCT
289 |  |  |  | 2008-07-02 |  | 080702_141215 | check firmware | VMEm2_6 |  | ME-2/2/36 | ALCT
255 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm2_6 |  | ME-2/2/36 | ALCT | Config problems for ALCT
275 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm2_6 |  | ME-2/2/36 |  | ALCT firmware incorrect
128 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm2_6 |  | ME-2/2/36 | ALCT | Config problems for ALCT: ME-2/2/36,
58 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm2_6 |  | ME-2/2/36 | ALCT: | Config problems for ALCT on ME-2/2/36, well known: blown fuse
594 |  |  |  | 2008-07-22 |  |  | CC | VMEm2_6 |  | ME-2/2/36, | ALCT: | Config problems for
471 |  |  |  | 2008-07-09 |  |  | CC | VMEm2_6 |  | ME-2/2/36, |  | Config problems for ALCT:
129 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm3_1 |  | ME-3/1/03 | DMB | Config problems for DMB: ME-3/1/3,
130 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm3_2 |  | ME-3/1/07 | DMB | Config problems for DMB: ME-3/1/7,
541 | 2 | access |  | 2008-07-18 |  |  | Misha | VMEm3_4 |  | ME-3/1/12 | CFEB1 | Hot SCA channels layer 1 CFEB1 - always there / Jason > does not look like a CFEB problem, investigate more | 90 | 2h | access chamber and swap CFEB | Y
203 |  |  |  | 2008-07-02 |  |  | Misha | VMEm3_4 |  | ME-3/1/12 |  | pl 1, input cable? |  | 2h | access to St 3
446 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_4 |  | ME-3/1/12 |  | Hot SCA channels layer 1 CFEB1 (see “other issues 1” plot) [low]
179 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_4 |  | ME-3/1/12 |  | ME-3/1/12 – Hot SCA channels layer 1 CFEB1 (see “other issues 1” plot)
204 | 3 | remove |  | 2008-07-18 |  |  | Misha | VMEm3_4 |  | ME-3/1/13 | CFEB3 | dead | 100 | 5h | swap >access, remove front chambers | Y
226 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_4 |  | ME-3/1/13 | CFEB3 | CFEB3 has no comparator data |  |  | check TMB cable, replace TMB or CFEB3
328 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_4 |  | ME-3/1/13 | CFEB3 | CFEB3 has no comparator data |  |  | check TMB cable, replace TMB or CFEB3
447 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_4 |  | ME-3/1/13 |  | no comparators data CFEB3 [medium]
78 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_4 |  | ME-3/1/13 | DFEB3 | – no comparators data CFEB3
21 |  |  |  | 2008-06-18 |  |  | Misha | VMEm3_4 |  | ME-3/1/13 | CFEB3 | CFEB3 has no comparator data |  |  | check TMB cable, replace TMB or CFEB3
180 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_4 |  | ME-3/1/13 |  | ME-3/1/13 – no comparators data CFEB5
636 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-3/1/15 |  | 4 cfeb2
548 | 5 |  |  | 2008-07-21 |  |  | Greg | VMEm3_5 |  | ME-3/1/15 | CFEB2 | kill_input=4
604 |  |  |  | 2008-07-20 |  |  | Misha | VMEm3_5 |  | ME-3/1/15 | CFEB#2 | channel link serial bit (lower bit) is broken
205 | 3 | remove |  | 2008-07-18 |  |  | Misha | VMEm3_5 |  | ME-3/1/15 | CFEB2 | dead channel link serial bit (middle bit) is broken | 90 | 5h | swap >access, remove front chambers | Y
375 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_5 |  | ME-3/1/15 | RUI#26, DDU Input 5: | CFEBs flagged with Bad CRCs: 0x02 = CFEB #2
376 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_5 |  | ME-3/1/15 | RUI#26, DDU Input 5: | CFEB CRC:80
389 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm3_5 |  | ME-3/1/15 | RUI#26, DDU Input 5: | CFEBs flagged with Bad CRCs: 0x02 = CFEB #2
390 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm3_5 |  | ME-3/1/15 | RUI#26, DDU Input 5: | CFEB CRC:456
227 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_5 |  | ME-3/1/15 | CFEB2 | CFEB2, bad Channel link bit |  |  | next: replace DMB or CFEB2
448 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_5 |  | ME-3/1/15 |  | 22% CFEB DAV Errors, 7% CFEB CRC and word count per sample errors, no CFEB2 SCA data [high]
79 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_5 |  | ME-3/1/15 | CFEB | - 22% CFEB DAV Errors, 7% CFEB CRC and word count per sample errors, no CFEB2 SCA data
98 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm3_5 |  | ME-3/1/15 | CFEB | - one “dead” CFEB?
22 |  |  |  | 2008-06-18 |  |  | Jason | VMEm3_5 |  | ME-3/1/15 | CFEB2 | CFEB2, bad Channel link bit |  |  | next: replace DMB or CFEB2
181 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_5 |  | ME-3/1/15 |  | 22 CFEB CRC and word count per sample errors, no CFEB2 SCA data
228 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_6 |  | ME-3/1/18 |  | needs timing scan, not done yet. DONE |  | done | timing tuned 23 June 08, ~6pm, fixed it.
426 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_6 |  | ME-3/1/18 |  | has only one event (timing?)
80 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_6 |  | ME-3/1/18 |  | #NAME?
99 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm3_6 |  | ME-3/1/18 | CFEB | - “bad”; CFEB problems? (wires look fine)
23 |  |  |  | 2008-06-18 |  |  | Misha | VMEm3_6 |  | ME-3/1/18 |  | needs timing scan, not done yet. DONE |  | done | timing tuned 23 June 08, ~6pm
182 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_6 |  | ME-3/1/18 |  | has only one event (timing?)
595 |  |  |  | 2008-07-22 |  |  | CC | VMEm3_2 |  | ME-3/1/7, | DMB: | Config problems for
453 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_6 |  | ME-3/2/01 |  | hot comparators channels in all layers border CFEB1-CFEB2 (Misha-> cable between CFEBs is missing)
82 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_6 |  | ME-3/2/01 | DFEB1 | – Hot comparator channels in all layers border CFEB1-CFEB2
83 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_6 |  | ME-3/2/01 | DFEB5 | – Dead comparator channels layers 1,4 CFEB5
84 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_6 |  | ME-3/2/01 | CFEB3 | - 25% CFEB DAV Errors, no CFEB3 SCA data
85 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_6 |  | ME-3/2/01 |  | - 80% CFEB, TMB, ALCT DAV errors and DMB Trailer missing errors
86 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_6 |  | ME-3/2/01 | CFEB3 | – Hot comparator channel layer 2,3 CFEB 3
183 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_6 |  | ME-3/2/01 |  | ME-3/2/01 – hot comparators channels in all layers border CFEB1-CFEB4
601 |  |  |  | 2008-07-20 |  |  | Misha |  |  | ME-3/2/03 | no HV, | The only missing chamber in run#51910 |  |  |  |  | XXX
542 | 4 | leave |  | 2008-07-18 |  |  | Misha | VMEm3_1 |  | ME-3/2/03 | HV | Turn off HV sector + disable interlock | 90 | 0h | external action only | Y
467 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEm3_1 |  | ME-3/2/03 |  | full chambers have HV OFF:
369 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_1 |  | ME-3/2/03 |  | Empty CSC list
335 |  |  |  | 2008-07-06 |  |  | Madorsky | VMEm3_1 |  | ME-3/2/03 |  | This chamber has a well-known problem (see my last elog), it is disconnected from HV until fixed. Please disable it in DCS.
344 |  |  |  | 2008-07-06 |  |  | Misha | VMEm3_1 |  | ME-3/2/03 |  | HV interlock problem - whole chamber;
441 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_1 |  | ME-3/2/03 |  | no data from CFEBs 3,4,5, very low frequency of reporting data (1st time with HV, needs timing in) [medium]
184 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_1 |  | ME-3/2/03 |  | ME-3/2/03 – (was masked/disabled) no data from CFEBs 3,4,5, very low frequency of reporting data
563 |  |  |  | 2008-07-22 |  |  | Luca | VMEm3_1 |  | ME-3/2/04 | ALCT | pushed check config button, and problem went away. Apparently it was the wrong configuration.
229 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_1 |  | ME-3/2/04 | CFEBs | bad slow-control readback for all CFEBs |  | done | replaced DMB 2008/Jun/11 13:00
24 |  |  |  | 2008-06-18 |  |  | Misha | VMEm3_1 |  | ME-3/2/04 | CFEBs | bad slow-control readback for all CFEBs |  |  | replaced DMB 2008/Jun/11 13:00
230 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_2 |  | ME-3/2/10 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors |  | done | timing tuned 23 June 08, ~6pm, fixed it.
427 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_2 |  | ME-3/2/10 |  | 98% events all 5 CFEBs report DMB Input FIFO full + B-words
81 |  |  |  | 2008-06-21 |  | 080621_154844 |  | VMEm3_2 |  | ME-3/2/10 | CFEB | - 98% events all 5 CFEBs report DMB Input FIFO full + B-words
100 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm3_2 |  | ME-3/2/10 | CFEB | - “bad”; CFEB problems? (wires look fine)
25 |  |  |  | 2008-06-18 |  |  | Jason | VMEm3_2 |  | ME-3/2/10 | CFEBs | all CFEBs have DAV-LCT and Full@DMB errors, GONE |  | done | timing tuned 23 June 08, ~6pm
185 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_2 |  | ME-3/2/10 |  | 98
564 |  |  |  | 2008-07-22 |  |  | Luca | VMEm3_2 |  | ME-3/2/12 | ALCT | , pushed config TMB+ALCT button, fixed the problem.
449 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_2 |  | ME-3/2/12 |  | HOT, comparators noise CFEB5, layers 1,4 (check TMB cable), SCA noise layer 1, 13% DAV errors, 7% CFEB CRC errors, dead anode channel layer 2 [high]
186 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_2 |  | ME-3/2/12 |  | HOT, comparators noise CFEB5, layers 1,4 (check TMB cable), SCA noise layer 1, 13 CFEB CRC errors, dead anode channel layer 2
256 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm3_3 |  | ME-3/2/17 | DMB | Config problems for DMB
276 |  |  |  | 2008-07-01 |  |  | Check firmware | VMEm3_3 |  | ME-3/2/17 | CFEB3 | CFEB firmware incorrect
131 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm3_3 |  | ME-3/2/17 | DMB | Config problems for DMB: ME-3/2/17,
231 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_3 |  | ME-3/2/19 | CFEB4 | CFEB4 hot CLCT |  |  | next: check TMB cable, replace TMB, DMB or CFEB4
26 |  |  |  | 2008-06-18 |  |  | Misha | VMEm3_3 |  | ME-3/2/19 | CFEB4 | CFEB4 hot CLCT |  |  | next: check TMB cable, replace TMB, DMB or CFEB4
209 |  |  |  | 2008-07-02 |  |  | Misha | VMEm3_4 |  | ME-3/2/21 | CFEB \1 & 2 | hot comparator |  |  | swap > access
637 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-3/2/22 |  | 5 cfeb3
546 |  |  |  | 2008-07-21 |  |  | Greg | VMEm3_4 |  | ME-3/2/22 | CFEB3 | kill_input=5
602 |  |  |  | 2008-07-20 |  |  | Misha | VMEm3_4 |  | ME-3/2/22 | CFEB#3 | channel link serial bit (middle bit) is broken
373 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_4 |  | ME-3/2/22 | RUI#24, DDU Input 3: | CFEBs flagged with Bad CRCs: 0x04 = CFEB #3
374 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_4 |  | ME-3/2/22 | RUI#24, DDU Input 3: | CFEB CRC:160
387 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm3_4 |  | ME-3/2/22 | RUI#24, DDU Input 3: | CFEBs flagged with Bad CRCs: 0x04 = CFEB #3
388 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm3_4 |  | ME-3/2/22 | RUI#24, DDU Input 3: | CFEB CRC:912
206 | 3 | access |  | 2008-07-02 |  |  | Misha | VMEm3_4 |  | ME-3/2/22 | CFEB3 | dead channel link serial bit (middle bit) is broken | 90 | 2h | swap >access | Y
232 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_4 |  | ME-3/2/22 | CFEB3 | CFEB3 bit 12 floating, does not follow cable |  |  | DMB swap done, still bad. next: replace CFEB3
450 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_4 |  | ME-3/2/22 |  | 25% CFEB DAV Errors caused no CFEB3 SCA data [high]
60 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm3_4 |  | ME-3/2/22 | DMB | Config problems for DMB: ME-3/2/22... went to DMB utils, clicked DMB CONTROL Load Firmware, Write Flash to Crate, Power-up Init.
101 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm3_4 |  | ME-3/2/22 | CFEB | - one “dead” CFEB + one “bad” CFEB?
27 |  |  |  | 2008-06-18 |  |  | Jason | VMEm3_4 |  | ME-3/2/22 | CFEB3 | CFEB3 bit 12 floating, does not follow cable |  |  | DMB swap done, no help. next: replace CFEB3
187 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_4 |  | ME-3/2/22 |  | 25
638 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-3/2/23 |  | 1 alct
565 |  |  |  | 2008-07-22 |  |  | Luca | VMEm3_4 |  | ME-3/2/23 | DMB | chip1 1.8 V reads as 1816 mV. Could not find anything else wrong.
549 |  |  |  | 2008-07-21 |  |  | Greg | VMEm3_4 |  | ME-3/2/23 | ALCT | kill_input=1
207 | 4 | remove |  | 2008-07-18 |  |  | Misha | VMEm3_4 |  | ME-3/2/23 | ALCT | always sends corrupted data - have to kill ALCT | 90 | 8h | remove front chambers | Y
466 |  |  |  | 2008-07-09 |  | CRUZET3 | Greg | VMEm3_4 |  | ME-3/2/23 |  | Bad ALCT/RAT
233 |  |  |  | 2008-07-02 |  |  | Jason | VMEm3_4 |  | ME-3/2/23 | ALCT | bad ALCT data, bit 15 often low in Header/Trailer, w/ 64-bit violations, shows as CFEB errors |  |  | replaced TMB 30 June 08 16:15, still bad. next: try RAT or ALCT
451 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_4 |  | ME-3/2/23 |  | 85% CFEB, TMB, ALCT DAV errors and DMB Trailer missing errors [high]
102 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm3_4 |  | ME-3/2/23 | CFEB | - “bad”/”dead”; CFEB problems? (wires look fine)
28 |  |  |  | 2008-06-18 |  |  | Jason | VMEm3_4 |  | ME-3/2/23 | ALCT | bad ALCT data, bit 15 often low in Header/Trailer, w/ 64-bit violations, shows as CFEB errors |  |  | next: check ALCT cables, replace TMB, RAT or ALCT
188 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_4 |  | ME-3/2/23 |  | 85
596 |  |  |  | 2008-07-22 |  |  | CC | VMEm3_4 |  | ME-3/2/23, | DMB: | Config problems for
566 |  |  |  | 2008-07-22 |  |  | Luca | VMEm3_4 |  | ME-3/2/24 | ALCT | slow control readings are wrong. Hard reset helped
452 |  |  |  | 2008-07-01 |  | 80630_131839 | Victor | VMEm3_4 |  | ME-3/2/24 |  | gap in comparators data for CFEB2 layers 2,3 (see “other issues 3” plot) [low]
105 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm3_4 |  | ME-3/2/24 |  | wires “hot”; (ME-1/2/17 is “masked”)
189 |  |  |  | 2008-06-30 |  | 080630_131839 | V. Barashko | VMEm3_4 |  | ME-3/2/24 |  | ME-3/2/24 – gap in comparators data for CFEB2 layers 2,3 (see “other issues 3” plot)
364 |  |  |  | 2008-07-06 |  |  | CC | VMEm3_5 |  | ME-3/2/27 | DMB: | Config problems for
472 |  |  |  | 2008-07-09 |  |  | CC | VMEm3_5 |  | ME-3/2/27, |  | Config problems for DMB:
377 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_5 |  | ME-3/2/29 | RUI#26, DDU Input 11: | CFEBs flagged with Bad CRCs: 0x10 = CFEB #5
378 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm3_5 |  | ME-3/2/29 | RUI#26, DDU Input 11: | CFEB CRC Error set 1 times
391 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm3_5 |  | ME-3/2/29 | RUI#26, DDU Input 11: | CFEBs flagged with Bad CRCs: 0x10 = CFEB #5
392 |  |  |  | 2008-07-07 |  | 50715 | Jason | VMEm3_5 |  | ME-3/2/29 | RUI#26, DDU Input 11: | CFEB CRC Error set 2 times
365 |  |  |  | 2008-07-06 |  |  | CC | VMEm3_6 |  | ME-3/2/34 | ALCT: | Config problems for
379 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm4_2 |  | ME-4/1/05 | RUI#33, DDU Input 6: | CFEBs flagged with Bad DAV-LCT: 0x01 = CFEB #1
380 |  |  |  | 2008-07-07 |  | 50709 | Jason | VMEm4_2 |  | ME-4/1/05 | RUI#33, DDU Input 6: | CFEB DAV-LCT match Error set 1 times
366 |  |  |  | 2008-07-06 |  |  | CC | VMEm4_3 |  | ME-4/1/08 | ALCT: | Config problems for
225 |  |  |  | 2008-07-02 |  |  | Jason | VMEm4_3 |  | ME-4/1/08 | CFEB1 | CFEB1 hot |  |  | timing tuned
20 |  |  |  | 2008-06-18 |  |  | Misha | VMEm4_3 |  | ME-4/1/08 | CFEB1 | CFEB1 hot |  |  | timing tuned
567 |  |  |  | 2008-07-22 |  |  | Luca | VMEm4_4 |  | ME-4/1/11 |  | dead yesterday, but is ok today. Neither power cycling nor hard resetting was able to bring this ALCT back yesterday. This problem needs to be resolved.
18 |  |  |  | 2008-06-18 |  |  | Misha | VMEm4_4 |  | ME-4/1/11 | ALCT | TMB slot 2 ALCT, bad communication |  | done | replaced TMB 6 June 08, 10am
223 |  |  |  | 2008-06-18 |  |  | Jason | VMEm4_4 |  | ME-4/1/11 | ALCT | TMB slot 2 ALCT, bad communication & configure |  |  | replaced TMB 6 June 08, 10am: still bad, try new RAT.
597 |  |  |  | 2008-07-22 |  |  | CC | VMEm4_4 |  | ME-4/1/11, | ALCT: | Config problems for
103 |  |  |  | 2008-06-21 |  | 080621_154844 | Stoyan Stoynev, Michael Schmitt, Andy Kubik NorthwesternUniversity | VMEm4_4 |  | ME-4/1/12 | CFEB | - “bad”? CFEB problems? (wires look fine)
208 |  |  |  | 2008-07-02 |  |  | Misha | VMEm4_5 |  | ME-4/1/14 | ALCT | dead
257 |  |  |  | 2008-07-01 |  | 20080701_171309 | Check Config | VMEm4_5 |  | ME-4/1/14 | ALCT | Config problems for ALCT
61 |  |  |  | 2008-06-24 | 11:00 |  | Check Config | VMEm4_5 |  | ME-4/1/14 | ALCT | Config problems for ALCT: ME-4/1/14, went away Configure TMB+ALCT fixed problem
19 |  |  |  | 2008-06-18 |  |  | Misha | VMEm4_5 |  | ME-4/1/14 | ALCT | TMB slot 2 ALCT on ME-4/1/14 – reads wrong thresholds for AFEBs 7-18 – configure of ALCT-TMB fix the problem each time from first attempt. |  |  | replaced TMB 25 June 08, 10am
224 |  |  |  | 2008-06-18 |  |  | Jason | VMEm4_5 |  | ME-4/1/14 | ALCT | TMB slot 2 ALCT on ME-4/1/14 – reads wrong thresholds for AFEBs 7-18 – configure problem of ALCT-TMB fixes the problem from first attempt. NOW ALCT Full@DMB and data Consistency error |  |  | replaced TMB 25 June 08, 10am. Replaced ALCT 06.30.08 16:00, now gives ALCT data errors. Do Phase/Timing Scan.
327 |  |  |  | 2008-06-18 |  |  | Jason | VMEm4_5 |  | ME-4/1/14 | ALCT | TMB slot 2 ALCT on ME-4/1/14 – reads wrong thresholds for AFEBs 7-18 – configure problem of ALCT-TMB fixes the problem from first attempt. NOW ALCT Full@DMB and data Consistency error |  |  | replaced TMB 25 June 08, 10am. Replaced ALCT 06.30.08 16:00, now gives ALCT data errors. Do Phase/Timing Scan.
639 |  |  |  | 2008-08-01 |  |  | Greg |  |  | ME-4/1/17 |  | kill at DDU timing checked to be OK
543 | 0 | OK |  | 2008-07-18 |  |  | Misha | VMEm4_5 |  | ME-4/1/17 | HV | one HV sector constant trip; - fixed |  |  |  | Y
343 |  |  |  | 2008-07-06 |  |  | Misha | VMEm4_5 |  | ME-4/1/17 |  | one HV sector constant trip;
132 |  |  |  | 2008-06-30 | 15:00 |  | Check Config | VMEm4_4 |  |  | CCB | Config problem for CCB/TTC | [5 empty] | XXX

## Input Worksheet
Entry # | Priority | Status / Action | Area | Date | Time | run # (During or Before) | Source | TS / Pcrate ID | Pcrate slot | chamber | On-chamber board | Problem Description | Confidence of Problem Solution | Estimated Repair Time | Comments on Action Needed | Chamber Summary | paste | VME | TEXT | STATUS | TEXT 2 | CHAMBER
PLUS SIDE: