## MPC Times
MPC Times
|  |  | Lengths in km |  |  | Measurement error is probably +- 2m
Maximum | [7 empty] | 0.087 |  |  |  |  | 0.103 |  |  |  |  | 0.11
Minimum | [7 empty] | 0.06 |  |  |  |  | 0.058 |  |  |  |  | 0.084
Disk |  |  | VME End | H |  |  |  | H |  | I |  |  | I |  | J |  |  | J
Station |  |  |  | YE-1 | [5 empty] | YE-2 |  |  |  |  | YE-3
Fiber | module | connector | Connector | Station | TS |  |  |  | ave | Station | TS |  |  | ave | Station | TS |  |  | ave
FIBER001 | 7 | 1 | M1 | 1 | 2 |  |  | 0.076 | 0.075375 | 3 | 1 |  | 0.091 | 0.091125 | 4 | 1 |  | 0.094 | 0.09275
FIBER002 | 7 | 2 | M1 | 1 | 3 |  |  | 0.076 |  | 2 | 1 |  | 0.093 |  | - | - |  | 0.092
FIBER003 | 7 | 3 | M2 | 1 | 2 |  |  | 0.075 |  | 3 | 1 |  | 0.092 |  | 4 | 1 |  | 0.09
FIBER004 | 7 | 4 | M2 | 1 | 3 |  |  | 0.075 |  | 2 | 1 |  | 0.089 |  | - | - |  | 0.093
FIBER005 | 7 | 5 | M3 | 1 | 2 |  |  | 0.076 |  | 3 | 1 |  | 0.093 |  | 4 | 1 |  | 0.094
FIBER006 | 7 | 6 | M3 | 1 | 3 |  |  | 0.075 |  | 2 | 1 |  | 0.091 |  | - | - |  | 0.091
FIBER007 | 7 | 7 | spare | 1 | 2 |  |  | 0.076 |  | 3 | 1 |  | 0.09 |  | 4 | 1 |  | 0.094
FIBER008 | 7 | 8 | spare | 1 | 3 |  |  | 0.074 |  | 2 | 1 |  | 0.09 |  | - | - |  | 0.094
FIBER009 | 8 | 1 | M1 | 1 | 5 |  |  | 0.081 | 0.084 | 3 | 2 | max | 0.103 | 0.099625 | 4 | 2 |  | 0.102 | 0.101
FIBER010 | 8 | 2 | M1 | 1 | 4 |  |  | 0.084 |  | 2 | 2 |  | 0.1 |  | - | - |  | 0.1
FIBER011 | 8 | 3 | M2 | 1 | 5 |  |  | 0.084 |  | 3 | 2 |  | 0.095 |  | 4 | 2 |  | 0.102
FIBER012 | 8 | 4 | M2 | 1 | 4 |  |  | 0.084 |  | 2 | 2 |  | 0.101 |  | - | - |  | 0.1
FIBER013 | 8 | 5 | M3 | 1 | 5 |  | max | 0.087 |  | 3 | 2 |  | 0.101 |  | 4 | 2 |  | 0.103
FIBER014 | 8 | 6 | M3 | 1 | 4 |  |  | 0.082 |  | 2 | 2 |  | 0.1 |  | - | - |  | 0.101
FIBER015 | 8 | 7 | spare | 1 | 5 |  |  | 0.085 |  | 3 | 2 |  | 0.097 |  | 4 | 2 |  | 0.1
FIBER016 | 8 | 8 | spare | 1 | 4 |  |  | 0.085 |  | 2 | 2 |  | 0.1 |  | - | - |  | 0.1
FIBER017 | 9 | 1 | M1 | 1 | 7 |  |  | 0.084 | 0.083 | 3 | 3 |  | 0.099 | 0.097429 | 4 | 3 |  | 0.105 | 0.10812
FIBER018 | 9 | 2 | M1 | 1 | 6 |  |  | 0.083 |  | 2 | 3 |  | 0.1 |  | - | - | max | 0.11
FIBER019 | 9 | 3 | M2 | 1 | 7 |  |  | 0.081 |  | 3 | 3 |  | 0.1 |  | 4 | 3 | max | 0.11
FIBER020 | 9 | 4 | M2 | 1 | 6 |  |  | 0.086 |  | 2 | 3 |  | 0.096 |  | - | - |  | 0.107
FIBER021 | 9 | 5 | M3 | 1 | 7 |  |  | 0.083 |  | 3 | 3 |  | 0.095 |  | 4 | 3 |  | 0.109
FIBER022 | 9 | 6 | M3 | 1 | 6 |  |  | 0.082 |  | 2 | 3 |  | 0.097 |  | - | - |  | 0.109
FIBER023 | 9 | 7 | spare | 1 | 7 |  |  | 0.083 |  | 3 | 3 |  | 0.095 |  | 4 | 3 |  | 0.106
FIBER024 | 9 | 8 | spare | 1 | 6 |  |  | 0.082 |  | 2 | 3 |  | 0.063 |  | - | - |  | 0.109
FIBER025 | 10 | 1 | M1 | 1 | 9 |  |  | 0.075 | 0.073625 | 2 | 4 |  | 0.058 | 0.092857 | 4 | 4 |  | 0.102 | 0.10312
FIBER026 | 10 | 2 | M1 | 1 | 8 |  |  | 0.072 |  | 3 | 4 |  | 0.093 |  | - | - |  | 0.102
FIBER027 | 10 | 3 | M2 | 1 | 9 |  |  | 0.071 |  | 2 | 4 |  | 0.094 |  | 4 | 4 |  | 0.104
FIBER028 | 10 | 4 | M2 | 1 | 8 |  |  | 0.074 |  | 3 | 4 |  | 0.093 |  | - | - |  | 0.103
FIBER029 | 10 | 5 | M3 | 1 | 9 |  |  | 0.075 |  | 2 | 4 |  | 0.093 |  | 4 | 4 |  | 0.102
FIBER030 | 10 | 6 | M3 | 1 | 8 |  |  | 0.075 |  | 3 | 4 |  | 0.092 |  | - | - |  | 0.106
FIBER031 | 10 | 7 | spare | 1 | 9 |  |  | 0.073 |  | 2 | 4 |  | 0.091 |  | 4 | 4 |  | 0.103
FIBER032 | 10 | 8 | spare | 1 | 8 |  |  | 0.074 |  | 3 | 4 |  | 0.094 |  | - | - |  | 0.103
FIBER033 | 11 | 1 | M1 | 1 | 10 |  |  | 0.061 | 0.061 | 3 | 5 |  | 0.082 | 0.08175 | 4 | 5 |  | 0.093 | 0.091875
FIBER034 | 11 | 2 | M1 | 1 | 11 |  |  | 0.061 |  | 2 | 5 |  | 0.081 |  | - | - |  | 0.092
FIBER035 | 11 | 3 | M2 | 1 | 10 |  |  | 0.06 |  | 3 | 5 |  | 0.082 |  | 4 | 5 |  | 0.091
FIBER036 | 11 | 4 | M2 | 1 | 11 |  |  | 0.062 |  | 2 | 5 |  | 0.082 |  | - | - |  | 0.092
FIBER037 | 11 | 5 | M3 | 1 | 10 |  |  | 0.061 |  | 3 | 5 |  | 0.084 |  | 4 | 5 |  | 0.092
FIBER038 | 11 | 6 | M3 | 1 | 11 |  |  | 0.061 |  | 2 | 5 |  | 0.081 |  | - | - |  | 0.09
FIBER039 | 11 | 7 | spare | 1 | 10 |  |  | 0.062 |  | 3 | 5 |  | 0.08 |  | 4 | 5 |  | 0.092
FIBER040 | 11 | 8 | spare | 1 | 11 |  |  | 0.06 |  | 2 | 5 |  | 0.082 |  | - | - |  | 0.093
FIBER041 | 12 | 1 | M1 | 1 | 12 |  |  | 0.064 | 0.067 | 3 | 6 |  | 0.082 | 0.082875 | 4 | 6 |  | 0.084 | 0.086125
FIBER042 | 12 | 2 | M1 | 1 | 1 |  |  | 0.065 |  | 2 | 6 |  | 0.083 |  | - | - |  | 0.088
FIBER043 | 12 | 3 | M2 | 1 | 12 |  |  | 0.069 |  | 3 | 6 |  | 0.084 |  | 4 | 6 |  | 0.087
FIBER044 | 12 | 4 | M2 | 1 | 1 |  |  | 0.068 |  | 2 | 6 |  | 0.083 |  | - | - |  | 0.085
FIBER045 | 12 | 5 | M3 | 1 | 12 |  |  | 0.07 |  | 3 | 6 |  | 0.082 |  | 4 | 6 |  | 0.086
FIBER046 | 12 | 6 | M3 | 1 | 1 |  |  | 0.068 |  | 2 | 6 |  | 0.081 |  | - | - |  | 0.087
FIBER047 | 12 | 7 | spare | 1 | 12 |  |  | 0.066 |  | 3 | 6 |  | 0.084 |  | 4 | 6 |  | 0.085
FIBER048 | 12 | 8 | spare | 1 | 1 |  |  | 0.066 |  | 2 | 6 |  | 0.084 |  | - | - |  | 0.087

## TTC Times
TTC Lengths
|  |  | Lengths in km |  |  | Measurement error is probably +- 1m
Maximum | [6 empty] | 0.09 |  |  |  | 0.108 |  |  |  | 0.116
Minimum | [6 empty] | 0.067 |  |  |  | 0.089 |  |  |  | 0
Disk |  |  | VME End | H |  |  | H | I |  |  | I | J |  |  | J
Station |  |  |  | YE-1 |  |  |  | YE-2 |  |  |  | YE-3
Fiber | module | connector | Connector | Station | TS |  |  | Station | TS |  |  | Station | TS
FIBER001 | 7 | 1 | 1 | 1 | 2 |  | 0.079 | 3 | 1 |  | 0.102 | 4 | 1 |  | 0.099
FIBER002 | 7 | 2 | 2 | 1 | 3 |  | 0.08 | 2 | 1 |  | 0.098 | - | - |  | 0.099
FIBER003 | 7 | 3 | 1 | 1 | 2 |  | 0.08 | 3 | 1 |  | 0.099 | 4 | 1 |  | 0.099
FIBER004 | 7 | 4 | 2 | 1 | 3 |  | 0.081 | 2 | 1 |  | 0.098 | - | - |  | 0.1
FIBER005 | 8 | 1 | 1 | 1 | 5 |  | 0.089 | 3 | 2 | max | 0.108 | 4 | 2 |  | 0.11
FIBER006 | 8 | 2 | 2 | 1 | 4 | max | 0.09 | 2 | 2 |  | 0.107 | - | - |  | 0.11
FIBER007 | 8 | 3 | 1 | 1 | 5 | max | 0.09 | 3 | 2 | max | 0.108 | 4 | 2 |  | 0.11
FIBER008 | 8 | 4 | 2 | 1 | 4 | max | 0.09 | 2 | 2 |  | 0.107 | - | - |  | 0.109
FIBER009 | 9 | 1 | 1 | 1 | 7 |  | 0.088 | 3 | 3 |  | 0.105 | 4 | 3 |  | 0.115
FIBER010 | 9 | 2 | 2 | 1 | 6 |  | 0.087 | 2 | 3 |  | 0.105 | - | - | max | 0.116
FIBER011 | 9 | 3 | 1 | 1 | 7 |  | 0.087 | 3 | 3 |  | 0.105 | 4 | 3 |  | 0.115
FIBER012 | 9 | 4 | 2 | 1 | 6 |  | 0.087 | 2 | 3 | max | 0.108 | - | - | max | 0.116
FIBER013 | 10 | 1 | 1 | 1 | 9 |  | 0.079 | 2 | 4 |  | 0.101 | 4 | 4 |  | 0.099
FIBER014 | 10 | 2 | 2 | 1 | 8 |  | 0.079 | 3 | 4 |  | 0.1 | - | - |  | 0.099
FIBER015 | 10 | 3 | 1 | 1 | 9 |  | 0.079 | 2 | 4 |  | 0.1 | 4 | 4 |  | 0.099
FIBER016 | 10 | 4 | 2 | 1 | 8 |  | 0.079 | 3 | 4 |  | 0.1 | - | - |  | 0.099
FIBER017 | 11 | 1 | 1 | 1 | 10 |  | 0.067 | 3 | 5 |  | 0.089 | 4 | 5 |  | 0.093
FIBER018 | 11 | 2 | 2 | 1 | 11 |  | 0.067 | 2 | 5 |  | 0.09 | - | - |  | 0.093
FIBER019 | 11 | 3 | 1 | 1 | 10 |  | 0.067 | 3 | 5 |  | 0.09 | 4 | 5 |  | 0.093
FIBER020 | 11 | 4 | 2 | 1 | 11 |  | 0.068 | 2 | 5 |  | 0.089 | - | - |  | 0.093
FIBER021 | 12 | 1 | 1 | 1 | 12 |  | 0.073 | 2 | 6 |  | 0.09 | 4 | 6 |  | 0
FIBER022 | 12 | 2 | 2 | 1 | 1 |  | 0.072 | 3 | 6 |  | 0.09 | - | - |  | 0
FIBER023 | 12 | 3 | 1 | 1 | 12 |  | 0.072 | 2 | 6 |  | 0.09 | 4 | 6 |  | 0
FIBER024 | 12 | 4 | 2 | 1 | 1 |  | 0.073 | 3 | 6 |  | 0.091 | - | - |  | 0

## S1 Rack Mapping
| S1D04h(YE1) |  | S1D04i(YE2) |  | S1D04j(YE3) |  | MPC-TF Fiber patch panel layout (USC55-S1D04)
Module # | a | b | a | b | a | b
1 | +1/2 | +1/3 | +3/1 | +2/1 | +4/1 | x | cassette:
2 | +1/5 | +1/4 | +3/2 | +2/2 | +4/2 | x | ----- |  | bot | top
3 | +1/7 | +1/6 | +3/3 | +2/3 | +4/3 | x | ¦a b¦ |  | M1 | M1
4 | +1/9 | +1/8 | +2/4 | +3/2 | +4/4 | x | ¦c d¦ |  | M2 | M2
5 | +1/10 | +1/11 | +3/5 | +2/5 | +4/5 | x | ¦e f¦ |  | M3 | M3
6 | +1/12 | +1/1 | +2/6 | +3/6 | +4/6 | x | ¦h i¦ |  | sp | sp
7 | -1/2 | -1/3 | -3/1 | -2/1 | -4/1 | x | ¦x x¦ |  | nc | nc
8 | -1/5 | -1/4 | -3/2 | -2/2 | -4/2 | x | ¦x x¦ |  | nc | nc
9 | -1/7 | -1/6 | -3/3 | -2/3 | -4/3 | x
10 | -1/9 | -1/8 | -2/4 | -3/4 | -4/4 | x
11 | -1/10 | -1/11 | -3/5 | -2/5 | -4/5 | x
12 | -1/12 | -1/1 | -2/6 | -3/6 | -4/6 | x
| S1D05j(YE1) |  | S1D05i(YE2) |  | S1D05h(YE3) |  | TTC-CCB Fiber patch panel layout (USC55-S1D05)
Module # | a | b | a | b | a | b
1 | +1/2 | +1/3 | +3/1 | +2/1 | +4/1 | x | cassette:
2 | +1/5 | +1/4 | +3/2 | +2/2 | +4/2 | x | ----- | bot | top
3 | +1/7 | +1/6 | +3/3 | +2/3 | +4/3 | x | ¦a b¦ | a | b | used
4 | +1/9 | +1/8 | +2/4 | +3/2 | +4/4 | x | ¦c d¦ | c | d | spares
5 | +1/10 | +1/11 | +3/5 | +2/5 | +4/5 | x | ¦x x¦ |  |  | not connected
6 | +1/12 | +1/1 | +2/6 | +3/6 | +4/6 | x | ¦x x¦
7 | -1/2 | -1/3 | -3/1 | -2/1 | -4/1 | x | ¦x x¦
8 | -1/5 | -1/4 | -3/2 | -2/2 | -4/2 | x | ¦x x¦
9 | -1/7 | -1/6 | -3/3 | -2/3 | -4/3 | x | ¦x x¦
10 | -1/9 | -1/8 | -2/4 | -3/4 | -4/4 | x
11 | -1/10 | -1/11 | -3/5 | -2/5 | -4/5 | x
12 | -1/12 | -1/1 | -2/6 | -3/6 | -4/6 | x

## combined
Combined Times | [8 empty] | Difference
Fiber | module | Rack | YE-1 MPC | YE-2 MPC | YE-3 MPC | YE-1 TTC | YE-2 TTC | YE-3 TTC | Rack | YE-1 | YE-2 | YE-3 | YE-1 | YE-2 | YE-3
FIBER001 | 7 | 1 | 0.076 | 0.091 | 0.094 | 0.079 | 0.102 | 0.099 | 1 | -3 | -10 | -6 | -4.625 | -8.125 | -6.5
FIBER002 | 7 | 1 | 0.076 | 0.093 | 0.092
FIBER003 | 7 | 1 | 0.075 | 0.092 | 0.09 | 0.08 | 0.098 | 0.099 | 1 | -5 | -7.5 | -7.5
FIBER004 | 7 | 1 | 0.075 | 0.089 | 0.093
FIBER005 | 7 | 1 | 0.076 | 0.093 | 0.094 | 0.08 | 0.099 | 0.099 | 1 | -4.5 | -7 | -6.5
FIBER006 | 7 | 1 | 0.075 | 0.091 | 0.091
FIBER007 | 7 | 1 | 0.076 | 0.09 | 0.094 | 0.081 | 0.098 | 0.1 | 1 | -6 | -8 | -6
FIBER008 | 7 | 1 | 0.074 | 0.09 | 0.094
FIBER009 | 8 | 2 | 0.081 | 0.103 | 0.102 | 0.089 | 0.108 | 0.11 | 2 | -6.5 | -6.5 | -9 | -5.75 | -7.875 | -8.75
FIBER010 | 8 | 2 | 0.084 | 0.1 | 0.1
FIBER011 | 8 | 2 | 0.084 | 0.095 | 0.102 | 0.09 | 0.107 | 0.11 | 2 | -6 | -9 | -9
FIBER012 | 8 | 2 | 0.084 | 0.101 | 0.1
FIBER013 | 8 | 2 | 0.087 | 0.101 | 0.103 | 0.09 | 0.108 | 0.11 | 2 | -5.5 | -7.5 | -8
FIBER014 | 8 | 2 | 0.082 | 0.1 | 0.101
FIBER015 | 8 | 2 | 0.085 | 0.097 | 0.1 | 0.09 | 0.107 | 0.109 | 2 | -5 | -8.5 | -9
FIBER016 | 8 | 2 | 0.085 | 0.1 | 0.1
FIBER017 | 9 | 3 | 0.084 | 0.099 | 0.105 | 0.088 | 0.105 | 0.115 | 3 | -4.5 | -5.5 | -7.5 | -4.25 | -12.625 | -7.375
FIBER018 | 9 | 3 | 0.083 | 0.1 | 0.11
FIBER019 | 9 | 3 | 0.081 | 0.1 | 0.11 | 0.087 | 0.105 | 0.116 | 3 | -3.5 | -7 | -7.5
FIBER020 | 9 | 3 | 0.086 | 0.096 | 0.107
FIBER021 | 9 | 3 | 0.083 | 0.095 | 0.109 | 0.087 | 0.105 | 0.115 | 3 | -4.5 | -9 | -6
FIBER022 | 9 | 3 | 0.082 | 0.097 | 0.109
FIBER023 | 9 | 3 | 0.083 | 0.095 | 0.106 | 0.087 | 0.108 | 0.116 | 3 | -4.5 | -29 | -8.5
FIBER024 | 9 | 3 | 0.082 | 0.063 | 0.109
FIBER025 | 10 | 4 | 0.075 | 0.058 | 0.102 | 0.079 | 0.101 | 0.099 | 4 | -5.5 | -25.5 | 3 | -5.375 | -11.75 | 4.125
FIBER026 | 10 | 4 | 0.072 | 0.093 | 0.102
FIBER027 | 10 | 4 | 0.071 | 0.094 | 0.104 | 0.079 | 0.1 | 0.099 | 4 | -6.5 | -6.5 | 4.5
FIBER028 | 10 | 4 | 0.074 | 0.093 | 0.103
FIBER029 | 10 | 4 | 0.075 | 0.093 | 0.102 | 0.079 | 0.1 | 0.099 | 4 | -4 | -7.5 | 5
FIBER030 | 10 | 4 | 0.075 | 0.092 | 0.106
FIBER031 | 10 | 4 | 0.073 | 0.091 | 0.103 | 0.079 | 0.1 | 0.099 | 4 | -5.5 | -7.5 | 4
FIBER032 | 10 | 4 | 0.074 | 0.094 | 0.103
FIBER033 | 11 | 5 | 0.061 | 0.082 | 0.093 | 0.067 | 0.089 | 0.093 | 5 | -6 | -7.5 | -0.5 | -6.25 | -7.75 | -1.125
FIBER034 | 11 | 5 | 0.061 | 0.081 | 0.092
FIBER035 | 11 | 5 | 0.06 | 0.082 | 0.091 | 0.067 | 0.09 | 0.093 | 5 | -6 | -8 | -1.5
FIBER036 | 11 | 5 | 0.062 | 0.082 | 0.092
FIBER037 | 11 | 5 | 0.061 | 0.084 | 0.092 | 0.067 | 0.09 | 0.093 | 5 | -6 | -7.5 | -2
FIBER038 | 11 | 5 | 0.061 | 0.081 | 0.09
FIBER039 | 11 | 5 | 0.062 | 0.08 | 0.092 | 0.068 | 0.089 | 0.093 | 5 | -7 | -8 | -0.5
FIBER040 | 11 | 5 | 0.06 | 0.082 | 0.093
FIBER041 | 12 | 6 | 0.064 | 0.082 | 0.084 | 0.073 | 0.09 | 0 | 6 | -8.5 | -7.5 | 86 | -5.5 | -7.375 | 86.125
FIBER042 | 12 | 6 | 0.065 | 0.083 | 0.088
FIBER043 | 12 | 6 | 0.069 | 0.084 | 0.087 | 0.072 | 0.09 | 0 | 6 | -3.5 | -6.5 | 86
FIBER044 | 12 | 6 | 0.068 | 0.083 | 0.085
FIBER045 | 12 | 6 | 0.07 | 0.082 | 0.086 | 0.072 | 0.09 | 0 | 6 | -3 | -8.5 | 86.5
FIBER046 | 12 | 6 | 0.068 | 0.081 | 0.087
FIBER047 | 12 | 6 | 0.066 | 0.084 | 0.085 | 0.073 | 0.091 | 0 | 6 | -7 | -7 | 86
FIBER048 | 12 | 6 | 0.066 | 0.084 | 0.087

## MPC-S1D04H
Job Information
Job ID: |  | Date: | 26.03.2008 10:06:36
Reason:
Contractor: |  | Operator A:
Customer: |  | Operator B:
Custom job 1: |  | Custom job 2:
Cable Information
Cable ID: | S1D04H
Location A: |  | Location B:
Cable Mfr.: |  | Cable Type:
Custom cable 1: |  | Custom cable 2:
System Information
| Unit A | Unit B
Unit Name: | FOT-932 | FOT-932
Serial Number: | 417390 | 417389
FasTesT
Fiber ID | Wavelength | Loss | Ref. | Loss | Ref. | Average | ORL | ORL | Length
|  | A->B | A->B | B->A | B->A |  | A | B
| (nm) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB)
(km)
FIBER001 | 1300 | 0.2 | 1.08 | 0.2 | 1.28 | 0.2 |  |  | 0.076
(26.03.2008 10:10:14)
FIBER002 | 1300 | 0.11 | 1.08 | 0.11 | 1.28 | 0.11 |  |  | 0.076
(26.03.2008 10:10:38)
FIBER003 | 1300 | 0.35 | 1.08 | 0.34 | 1.28 | 0.35 |  |  | 0.075
(26.03.2008 10:11:07)
FIBER004 | 1300 | 0.22 | 1.08 | 0.22 | 1.28 | 0.22 |  |  | 0.075
(26.03.2008 10:11:28)
FIBER005 | 1300 | 0.25 | 1.08 | 0.25 | 1.28 | 0.25 |  |  | 0.076
(26.03.2008 10:11:48)
FIBER006 | 1300 | 0.19 | 1.08 | 0.2 | 1.28 | 0.19 |  |  | 0.075
(26.03.2008 10:12:07)
FIBER007 | 1300 | 0.48 | 1.08 | 0.48 | 1.28 | 0.48 |  |  | 0.076
(26.03.2008 10:12:28)
FIBER008 | 1300 | 0.26 | 1.08 | 0.25 | 1.28 | 0.26 |  |  | 0.074
(26.03.2008 10:12:43)
FIBER009 | 1300 | 0.37 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.081
(26.03.2008 10:13:47)
FIBER010 | 1300 | 0.44 | 1.08 | 0.44 | 1.28 | 0.44 |  |  | 0.084
(26.03.2008 10:14:10)
FIBER011 | 1300 | 0.27 | 1.08 | 0.27 | 1.28 | 0.27 |  |  | 0.084
(26.03.2008 10:14:30)
FIBER012 | 1300 | 0.43 | 1.08 | 0.43 | 1.28 | 0.43 |  |  | 0.084
(26.03.2008 10:18:13)
FIBER013 | 1300 | 0.53 | 1.08 | 0.52 | 1.28 | 0.53 |  |  | 0.087
(26.03.2008 10:18:33)
FIBER014 | 1300 | 0.76 | 1.08 | 0.75 | 1.28 | 0.75 |  |  | 0.082
(26.03.2008 10:18:58)
FIBER015 | 1300 | 0.56 | 1.08 | 0.56 | 1.28 | 0.56 |  |  | 0.085
(26.03.2008 10:19:22)
FIBER016 | 1300 | 0.56 | 1.08 | 0.55 | 1.28 | 0.55 |  |  | 0.085
(26.03.2008 10:19:38)
FIBER017 | 1300 | 0.52 | 1.08 | 0.52 | 1.28 | 0.52 |  |  | 0.084
(26.03.2008 10:22:54)
FIBER018 | 1300 | 0.36 | 1.08 | 0.35 | 1.28 | 0.35 |  |  | 0.083
(26.03.2008 10:23:11)
FIBER019 | 1300 | 0.38 | 1.08 | 0.37 | 1.28 | 0.38 |  |  | 0.081
(26.03.2008 10:23:40)
FIBER020 | 1300 | 0.28 | 1.08 | 0.28 | 1.28 | 0.28 |  |  | 0.086
(26.03.2008 10:24:02)
FIBER021 | 1300 | 0.39 | 1.08 | 0.39 | 1.28 | 0.39 |  |  | 0.083
(26.03.2008 10:24:22)
FIBER022 | 1300 | 0.38 | 1.08 | 0.37 | 1.28 | 0.38 |  |  | 0.082
(26.03.2008 10:24:40)
FIBER023 | 1300 | 0.47 | 1.08 | 0.46 | 1.28 | 0.47 |  |  | 0.083
(26.03.2008 10:25:00)
FIBER024 | 1300 | 0.42 | 1.08 | 0.42 | 1.28 | 0.42 |  |  | 0.082
(26.03.2008 10:25:16)
FIBER025 | 1300 | 0.17 | 1.08 | 0.17 | 1.28 | 0.17 |  |  | 0.075
(26.03.2008 10:27:25)
FIBER026 | 1300 | 0.3 | 1.08 | 0.3 | 1.28 | 0.3 |  |  | 0.072
(26.03.2008 10:27:41)
FIBER027 | 1300 | 0.66 | 1.08 | 0.66 | 1.28 | 0.66 |  |  | 0.071
(26.03.2008 10:28:02)
FIBER028 | 1300 | 0.33 | 1.08 | 0.32 | 1.28 | 0.33 |  |  | 0.074
(26.03.2008 10:28:21)
FIBER029 | 1300 | 0.15 | 1.08 | 0.15 | 1.28 | 0.15 |  |  | 0.075
(26.03.2008 10:28:40)
FIBER030 | 1300 | 0.33 | 1.08 | 0.33 | 1.28 | 0.33 |  |  | 0.075
(26.03.2008 10:29:00)
FIBER031 | 1300 | 0.44 | 1.08 | 0.44 | 1.28 | 0.44 |  |  | 0.073
(26.03.2008 10:29:23)
FIBER032 | 1300 | 0.39 | 1.08 | 0.39 | 1.28 | 0.39 |  |  | 0.074
(26.03.2008 10:29:40)
FIBER033 | 1300 | 0.23 | 1.08 | 0.22 | 1.28 | 0.23 |  |  | 0.061
(26.03.2008 10:32:40)
FIBER034 | 1300 | 0.24 | 1.08 | 0.24 | 1.28 | 0.24 |  |  | 0.061
(26.03.2008 10:32:55)
FIBER035 | 1300 | 0.31 | 1.08 | 0.3 | 1.28 | 0.3 |  |  | 0.06
(26.03.2008 10:33:15)
FIBER036 | 1300 | 0.17 | 1.08 | 0.17 | 1.28 | 0.17 |  |  | 0.062
(26.03.2008 10:33:38)
FIBER037 | 1300 | 0.39 | 1.08 | 0.38 | 1.28 | 0.38 |  |  | 0.061
(26.03.2008 10:34:02)
FIBER038 | 1300 | 0.39 | 1.08 | 0.38 | 1.28 | 0.39 |  |  | 0.061
(26.03.2008 10:34:18)
FIBER039 | 1300 | 0.78 | 1.08 | 0.78 | 1.28 | 0.78 |  |  | 0.062
(26.03.2008 10:34:47)
FIBER040 | 1300 | 0.57 | 1.08 | 0.56 | 1.28 | 0.57 |  |  | 0.06
(26.03.2008 10:35:06)
FIBER041 | 1300 | 0.59 | 1.08 | 0.59 | 1.28 | 0.59 |  |  | 0.064
(26.03.2008 10:41:37)
FIBER042 | 1300 | 0.3 | 1.08 | 0.3 | 1.28 | 0.3 |  |  | 0.065
(26.03.2008 10:41:55)
FIBER043 | 1300 | 0.18 | 1.08 | 0.18 | 1.28 | 0.18 |  |  | 0.069
(26.03.2008 10:42:14)
FIBER044 | 1300 | 0.37 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.068
(26.03.2008 10:42:31)
FIBER045 | 1300 | 0.2 | 1.08 | 0.21 | 1.28 | 0.2 |  |  | 0.07
(26.03.2008 10:42:53)
FIBER046 | 1300 | 0.16 | 1.08 | 0.16 | 1.28 | 0.16 |  |  | 0.068
(26.03.2008 10:43:17)
FIBER047 | 1300 | 0.5 | 1.08 | 0.5 | 1.28 | 0.5 |  |  | 0.066
(26.03.2008 10:43:35)
FIBER048 | 1300 | 0.39 | 1.08 | 0.39 | 1.28 | 0.39 |  |  | 0.066
(26.03.2008 10:43:52)

## MPC-S1D04I
Job Information
Job ID: |  | Date: | 26.03.2008 10:45:33
Reason:
Contractor: |  | Operator A:
Customer: |  | Operator B:
Custom job 1: |  | Custom job 2:
Cable Information
Cable ID: | S1D04I
Location A: |  | Location B:
Cable Mfr.: |  | Cable Type:
Custom cable 1: |  | Custom cable 2:
System Information
| Unit A | Unit B
Unit Name: | FOT-932 | FOT-932
Serial Number: | 417390 | 417389
FasTesT
Fiber ID | Wavelength | Loss | Ref. | Loss | Ref. | Average | ORL | ORL | Length
|  | A->B | A->B | B->A | B->A |  | A | B
| (nm) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB)
(km)
FIBER001 | 1300 | 0.2 | 1.08 | 0.21 | 1.28 | 0.2 |  |  | 0.091
(26.03.2008 11:09:17)
FIBER002 | 1300 | 0.02 | 1.08 | 0.03 | 1.28 | 0.03 |  |  | 0.093
(26.03.2008 11:09:35)
FIBER003 | 1300 | 0.31 | 1.08 | 0.32 | 1.28 | 0.31 |  |  | 0.092
(26.03.2008 11:09:48)
FIBER004 | 1300 | 0.35 | 1.08 | 0.37 | 1.28 | 0.36 |  |  | 0.089
(26.03.2008 11:10:07)
FIBER005 | 1300 | 0.79 | 1.08 | 0.8 | 1.28 | 0.8 |  |  | 0.093
(26.03.2008 11:10:21)
FIBER006 | 1300 | 0.34 | 1.08 | 0.35 | 1.28 | 0.35 |  |  | 0.091
(26.03.2008 11:10:35)
FIBER007 | 1300 | 0.43 | 1.08 | 0.44 | 1.28 | 0.43 |  |  | 0.09
(26.03.2008 11:10:48)
FIBER008 | 1300 | 0.34 | 1.08 | 0.35 | 1.28 | 0.35 |  |  | 0.09
(26.03.2008 11:11:03)
FIBER009 | 1300 | 0.51 | 1.08 | 0.52 | 1.28 | 0.51 |  |  | 0.103
(26.03.2008 11:11:50)
FIBER010 | 1300 | 0.21 | 1.08 | 0.22 | 1.28 | 0.21 |  |  | 0.1
(26.03.2008 11:12:04)
FIBER011 | 1300 | 0.7 | 1.08 | 0.71 | 1.28 | 0.71 |  |  | 0.095
(26.03.2008 11:12:17)
FIBER012 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.36 |  |  | 0.101
(26.03.2008 11:12:33)
FIBER013 | 1300 | 0.47 | 1.08 | 0.48 | 1.28 | 0.47 |  |  | 0.101
(26.03.2008 11:12:47)
FIBER014 | 1300 | 0.25 | 1.08 | 0.26 | 1.28 | 0.25 |  |  | 0.1
(26.03.2008 11:13:01)
FIBER015 | 1300 | 0.49 | 1.08 | 0.5 | 1.28 | 0.5 |  |  | 0.097
(26.03.2008 11:13:16)
FIBER016 | 1300 | 0.37 | 1.08 | 0.38 | 1.28 | 0.37 |  |  | 0.1
(26.03.2008 11:13:30)
FIBER017 | 1300 | 0.21 | 1.08 | 0.23 | 1.28 | 0.22 |  |  | 0.099
(26.03.2008 11:15:16)
FIBER018 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.1
(26.03.2008 11:15:31)
FIBER019 | 1300 | 0.25 | 1.08 | 0.26 | 1.28 | 0.25 |  |  | 0.1
(26.03.2008 11:15:45)
FIBER020 | 1300 | 0.27 | 1.08 | 0.28 | 1.28 | 0.28 |  |  | 0.096
(26.03.2008 11:16:00)
FIBER021 | 1300 | 0.39 | 1.08 | 0.39 | 1.28 | 0.39 |  |  | 0.095
(26.03.2008 11:16:15)
FIBER022 | 1300 | 0.57 | 1.08 | 0.58 | 1.28 | 0.57 |  |  | 0.097
(26.03.2008 11:16:29)
FIBER023 | 1300 | 0.39 | 1.08 | 0.4 | 1.28 | 0.39 |  |  | 0.095
(26.03.2008 11:16:43)
FIBER024 | 1300 | 0.87 | 1.08 | 0.97 | 1.28 | 0.92 |  |  | 0.063
(26.03.2008 11:16:57)
FIBER025 | 1300 | 0.8 | 1.08 | 0.89 | 1.28 | 0.85 |  |  | 0.058
(26.03.2008 11:19:00)
FIBER026 | 1300 | 0.69 | 1.08 | 0.7 | 1.28 | 0.7 |  |  | 0.093
(26.03.2008 11:19:15)
FIBER027 | 1300 | 0.35 | 1.08 | 0.36 | 1.28 | 0.36 |  |  | 0.094
(26.03.2008 11:19:34)
FIBER028 | 1300 | 0.16 | 1.08 | 0.17 | 1.28 | 0.17 |  |  | 0.093
(26.03.2008 11:19:50)
FIBER029 | 1300 | 0.39 | 1.08 | 0.4 | 1.28 | 0.4 |  |  | 0.093
(26.03.2008 11:20:04)
FIBER030 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.092
(26.03.2008 11:20:18)
FIBER031 | 1300 | 0.31 | 1.08 | 0.32 | 1.28 | 0.32 |  |  | 0.091
(26.03.2008 11:20:32)
FIBER032 | 1300 | 0.56 | 1.08 | 0.56 | 1.28 | 0.56 |  |  | 0.094
(26.03.2008 11:20:47)
FIBER033 | 1300 | 0.26 | 1.08 | 0.27 | 1.28 | 0.27 |  |  | 0.082
(26.03.2008 11:22:55)
FIBER034 | 1300 | 0.28 | 1.08 | 0.29 | 1.28 | 0.28 |  |  | 0.081
(26.03.2008 11:23:15)
FIBER035 | 1300 | 0.33 | 1.08 | 0.34 | 1.28 | 0.33 |  |  | 0.082
(26.03.2008 11:23:30)
FIBER036 | 1300 | 0.38 | 1.08 | 0.39 | 1.28 | 0.39 |  |  | 0.082
(26.03.2008 11:23:52)
FIBER037 | 1300 | 0.28 | 1.08 | 0.29 | 1.28 | 0.29 |  |  | 0.084
(26.03.2008 11:24:06)
FIBER038 | 1300 | 0.37 | 1.08 | 0.38 | 1.28 | 0.37 |  |  | 0.081
(26.03.2008 11:24:22)
FIBER039 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.08
(26.03.2008 11:24:36)
FIBER040 | 1300 | 0.37 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.082
(26.03.2008 11:24:52)
FIBER041 | 1300 | 0.5 | 1.08 | 0.5 | 1.28 | 0.5 |  |  | 0.082
(26.03.2008 11:26:37)
FIBER042 | 1300 | 0.32 | 1.08 | 0.32 | 1.28 | 0.32 |  |  | 0.083
(26.03.2008 11:26:51)
FIBER043 | 1300 | 0.23 | 1.08 | 0.24 | 1.28 | 0.24 |  |  | 0.084
(26.03.2008 11:27:05)
FIBER044 | 1300 | 0.47 | 1.08 | 0.48 | 1.28 | 0.48 |  |  | 0.083
(26.03.2008 11:27:19)
FIBER045 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.36 |  |  | 0.082
(26.03.2008 11:27:35)
FIBER046 | 1300 | 0.33 | 1.08 | 0.33 | 1.28 | 0.33 |  |  | 0.081
(26.03.2008 11:27:53)
FIBER047 | 1300 | 0.35 | 1.08 | 0.36 | 1.28 | 0.36 |  |  | 0.084
(26.03.2008 11:28:08)
FIBER048 | 1300 | 0.4 | 1.08 | 0.41 | 1.28 | 0.41 |  |  | 0.084
(26.03.2008 11:28:21)

## MPC-S1D04J
Job Information
Job ID: |  | Date: | 26.03.2008 11:30:09
Reason:
Contractor: |  | Operator A:
Customer: |  | Operator B:
Custom job 1: |  | Custom job 2:
Cable Information
Cable ID: | S1D04J
Location A: |  | Location B:
Cable Mfr.: |  | Cable Type:
Custom cable 1: |  | Custom cable 2:
System Information
| Unit A | Unit B
Unit Name: | FOT-932 | FOT-932
Serial Number: | 417390 | 417389
FasTesT
Fiber ID | Wavelength | Loss | Ref. | Loss | Ref. | Average | ORL | ORL | Length
|  | A->B | A->B | B->A | B->A |  | A | B
| (nm) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB)
(km)
FIBER001 | 1300 | 0.43 | 1.08 | 0.44 | 1.28 | 0.43 |  |  | 0.094
(26.03.2008 11:30:20)
FIBER002 | 1300 | 0.28 | 1.08 | 0.28 | 1.28 | 0.28 |  |  | 0.092
(26.03.2008 11:30:35)
FIBER003 | 1300 | 0.2 | 1.08 | 0.21 | 1.28 | 0.21 |  |  | 0.09
(26.03.2008 11:30:50)
FIBER004 | 1300 | 0.28 | 1.08 | 0.28 | 1.28 | 0.28 |  |  | 0.093
(26.03.2008 11:31:05)
FIBER005 | 1300 | 0.21 | 1.08 | 0.22 | 1.28 | 0.21 |  |  | 0.094
(26.03.2008 11:31:18)
FIBER006 | 1300 | 0.06 | 1.08 | 0.07 | 1.28 | 0.06 |  |  | 0.091
(26.03.2008 11:31:33)
FIBER007 | 1300 | 0.34 | 1.08 | 0.34 | 1.28 | 0.34 |  |  | 0.094
(26.03.2008 11:31:48)
FIBER008 | 1300 | 0.7 | 1.08 | 0.71 | 1.28 | 0.71 |  |  | 0.094
(26.03.2008 11:32:04)
FIBER009 | 1300 | 0.67 | 1.08 | 0.68 | 1.28 | 0.67 |  |  | 0.102
(26.03.2008 11:36:50)
FIBER010 | 1300 | 0.35 | 1.08 | 0.36 | 1.28 | 0.35 |  |  | 0.1
(26.03.2008 11:37:05)
FIBER011 | 1300 | 0.24 | 1.08 | 0.24 | 1.28 | 0.24 |  |  | 0.102
(26.03.2008 11:37:20)
FIBER012 | 1300 | 0.53 | 1.08 | 0.53 | 1.28 | 0.53 |  |  | 0.1
(26.03.2008 11:37:34)
FIBER013 | 1300 | 0.43 | 1.08 | 0.43 | 1.28 | 0.43 |  |  | 0.103
(26.03.2008 11:37:48)
FIBER014 | 1300 | 0.34 | 1.08 | 0.35 | 1.28 | 0.35 |  |  | 0.101
(26.03.2008 11:38:02)
FIBER015 | 1300 | 0.49 | 1.08 | 0.5 | 1.28 | 0.5 |  |  | 0.1
(26.03.2008 11:39:31)
FIBER016 | 1300 | 0.37 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.1
(26.03.2008 11:39:52)
FIBER017 | 1300 | 0.38 | 1.08 | 0.38 | 1.28 | 0.38 |  |  | 0.105
(26.03.2008 11:41:33)
FIBER018 | 1300 | 0.43 | 1.08 | 0.43 | 1.28 | 0.43 |  |  | 0.11
(26.03.2008 11:41:48)
FIBER019 | 1300 | 0.08 | 1.08 | 0.09 | 1.28 | 0.09 |  |  | 0.11
(26.03.2008 11:42:02)
FIBER020 | 1300 | 0.15 | 1.08 | 0.15 | 1.28 | 0.15 |  |  | 0.107
(26.03.2008 11:42:16)
FIBER021 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.109
(26.03.2008 11:42:31)
FIBER022 | 1300 | 0.26 | 1.08 | 0.27 | 1.28 | 0.26 |  |  | 0.109
(26.03.2008 11:42:45)
FIBER023 | 1300 | 0.33 | 1.08 | 0.34 | 1.28 | 0.34 |  |  | 0.106
(26.03.2008 11:42:59)
FIBER024 | 1300 | 0.45 | 1.08 | 0.46 | 1.28 | 0.46 |  |  | 0.109
(26.03.2008 11:43:13)
FIBER025 | 1300 | 0.51 | 1.08 | 0.51 | 1.28 | 0.51 |  |  | 0.102
(26.03.2008 11:44:52)
FIBER026 | 1300 | 0.36 | 1.08 | 0.37 | 1.28 | 0.36 |  |  | 0.102
(26.03.2008 11:45:06)
FIBER027 | 1300 | 0.36 | 1.08 | 0.36 | 1.28 | 0.36 |  |  | 0.104
(26.03.2008 11:45:20)
FIBER028 | 1300 | 0.41 | 1.08 | 0.41 | 1.28 | 0.41 |  |  | 0.103
(26.03.2008 11:45:38)
FIBER029 | 1300 | 0.32 | 1.08 | 0.33 | 1.28 | 0.33 |  |  | 0.102
(26.03.2008 11:45:51)
FIBER030 | 1300 | 0.37 | 1.08 | 0.37 | 1.28 | 0.37 |  |  | 0.106
(26.03.2008 11:46:05)
FIBER031 | 1300 | 0.32 | 1.08 | 0.32 | 1.28 | 0.32 |  |  | 0.103
(26.03.2008 11:46:19)
FIBER032 | 1300 | 0.34 | 1.08 | 0.34 | 1.28 | 0.34 |  |  | 0.103
(26.03.2008 11:46:34)
FIBER033 | 1300 | 0.08 | 1.08 | 0.08 | 1.28 | 0.08 |  |  | 0.093
(26.03.2008 11:47:37)
FIBER034 | 1300 | 0.28 | 1.08 | 0.29 | 1.28 | 0.28 |  |  | 0.092
(26.03.2008 11:47:52)
FIBER035 | 1300 | 0.31 | 1.08 | 0.31 | 1.28 | 0.31 |  |  | 0.091
(26.03.2008 11:48:13)
FIBER036 | 1300 | 0.38 | 1.08 | 0.38 | 1.28 | 0.38 |  |  | 0.092
(26.03.2008 11:48:29)
FIBER037 | 1300 | 0.24 | 1.08 | 0.24 | 1.28 | 0.24 |  |  | 0.092
(26.03.2008 11:48:44)
FIBER038 | 1300 | 0.24 | 1.08 | 0.24 | 1.28 | 0.24 |  |  | 0.09
(26.03.2008 11:49:01)
FIBER039 | 1300 | 0.32 | 1.08 | 0.32 | 1.28 | 0.32 |  |  | 0.092
(26.03.2008 11:49:15)
FIBER040 | 1300 | 0.3 | 1.08 | 0.3 | 1.28 | 0.3 |  |  | 0.093
(26.03.2008 11:49:29)
FIBER041 | 1300 | 0.29 | 1.08 | 0.3 | 1.28 | 0.3 |  |  | 0.084
(26.03.2008 11:50:50)
FIBER042 | 1300 | 0.33 | 1.08 | 0.34 | 1.28 | 0.34 |  |  | 0.088
(26.03.2008 11:51:03)
FIBER043 | 1300 | 0.42 | 1.08 | 0.42 | 1.28 | 0.42 |  |  | 0.087
(26.03.2008 11:51:17)
FIBER044 | 1300 | 0.39 | 1.08 | 0.4 | 1.28 | 0.4 |  |  | 0.085
(26.03.2008 11:51:32)
FIBER045 | 1300 | 0.23 | 1.08 | 0.23 | 1.28 | 0.23 |  |  | 0.086
(26.03.2008 11:51:46)
FIBER046 | 1300 | 0.2 | 1.08 | 0.2 | 1.28 | 0.2 |  |  | 0.087
(26.03.2008 11:52:02)
FIBER047 | 1300 | 0.33 | 1.08 | 0.33 | 1.28 | 0.33 |  |  | 0.085
(26.03.2008 11:52:16)
FIBER048 | 1300 | 0.3 | 1.08 | 0.3 | 1.28 | 0.3 |  |  | 0.087
(26.03.2008 11:52:30)

## TTC-S1D05H
Job Information
Job ID: |  | Date: | 25.03.2008 15:04:03
Reason:
Contractor: |  | Operator A:
Customer: |  | Operator B:
Custom job 1: |  | Custom job 2:
Cable Information
Cable ID: | S1D05H
Location A: |  | Location B:
Cable Mfr.: |  | Cable Type:
Custom cable 1: |  | Custom cable 2:
System Information
| Unit A | Unit B
Unit Name: | FOT-932 | FOT-932
Serial Number: | 417389 | 417390
FasTesT
Fiber ID | Wavelength | Loss | Ref. | Loss | Ref. | Average | ORL | ORL | Length
|  | A->B | A->B | B->A | B->A |  | A | B
| (nm) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB)
(km)
FIBER001 | 1310 | 0.92 | 0.3 | 0.9 | -0.68 | 0.91 | 9.08 | 9.53 | 0.079
| 1550 | 0.55 | 0.26 | 0.47 | -0.36 | 0.51 | 13.8 | 14.16
(25.03.2008 15:26:34)
FIBER002 | 1310 | 0.73 | 0.3 | 0.9 | -0.68 | 0.81 | 44.46 | 43.96 | 0.08
| 1550 | 0.59 | 0.26 | 0.58 | -0.36 | 0.58 | 44.8 | 44.95
(25.03.2008 15:37:06)
FIBER003 | 1310 | 0.19 | 0.3 | 0.17 | -0.68 | 0.18 | 39.3 | 40.83 | 0.08
| 1550 | 0.15 | 0.26 | 0.18 | -0.36 | 0.16 | 39.54 | 40.59
(25.03.2008 15:37:24)
FIBER004 | 1310 | 0.39 | 0.3 | 0.4 | -0.68 | 0.39 | 38.38 | 38.56 | 0.081
| 1550 | 0.37 | 0.26 | 0.27 | -0.36 | 0.32 | 38.71 | 38.69
(25.03.2008 15:37:42)
FIBER005 | 1310 | 0.33 | 0.3 | 0.34 | -0.68 | 0.33 | 42.21 | 43.19 | 0.089
| 1550 | 0.36 | 0.26 | 0.31 | -0.36 | 0.34 | 43.07 | 43.92
(25.03.2008 15:47:26)
FIBER006 | 1310 | 0.31 | 0.3 | 0.33 | -0.68 | 0.32 | 42.15 | 43.06 | 0.09
| 1550 | 0.34 | 0.26 | 0.3 | -0.36 | 0.32 | 42.89 | 43.73
(25.03.2008 15:47:41)
FIBER007 | 1310 | 1.38 | 0.3 | 0.92 | -0.68 | 1.14 | 18.6 | 19.95 | 0.09
| 1550 | 1.14 | 0.26 | 0.75 | -0.36 | 0.94 | 21.51 | 24.09
(25.03.2008 15:48:05)
FIBER008 | 1310 | 0.19 | 0.3 | 0.17 | -0.68 | 0.18 | 43.71 | >43.89 | 0.09
| 1550 | 0.24 | 0.26 | 0.15 | -0.36 | 0.19 | 44.51 | >44.59
(25.03.2008 15:48:30)
FIBER009 | 1310 | 0.32 | 0.3 | 0.33 | -0.68 | 0.33 | >44.20 | >44.19 | 0.088
| 1550 | 0.31 | 0.26 | 0.19 | -0.36 | 0.25 | 45.07 | >44.67
(25.03.2008 15:50:48)
FIBER010 | 1310 | 0.34 | 0.3 | 0.25 | -0.68 | 0.3 | 43.87 | >44.03 | 0.087
| 1550 | 0.33 | 0.26 | 0.1 | -0.36 | 0.22 | 44.64 | >44.50
(25.03.2008 15:51:06)
FIBER011 | 1310 | 0.45 | 0.3 | 0.46 | -0.68 | 0.45 | 43.99 | >44.41 | 0.087
| 1550 | 0.46 | 0.26 | 0.41 | -0.36 | 0.44 | 44.91 | >45.07
(25.03.2008 15:51:24)
FIBER012 | 1310 | 0.27 | 0.3 | 0.29 | -0.68 | 0.28 | >44.10 | >44.11 | 0.087
| 1550 | 0.29 | 0.26 | 0.18 | -0.36 | 0.23 | >45.41 | >44.65
(25.03.2008 15:51:43)
FIBER013 | 1310 | 0.23 | 0.3 | 0.23 | -0.68 | 0.23 | >44.03 | >43.99 | 0.079
| 1550 | 0.26 | 0.26 | 0.06 | -0.36 | 0.15 | 45 | >44.42
(25.03.2008 15:53:55)
FIBER014 | 1310 | 0.14 | 0.3 | 0.12 | -0.68 | 0.13 | >43.86 | >43.79 | 0.079
| 1550 | 0.16 | 0.26 | 0.09 | -0.36 | 0.13 | >45.18 | >44.47
(25.03.2008 15:54:25)
FIBER015 | 1310 | 0.29 | 0.3 | 0.3 | -0.68 | 0.29 | 43.12 | 44 | 0.079
| 1550 | 0.2 | 0.26 | 0.2 | -0.36 | 0.2 | 44.23 | >44.68
(25.03.2008 15:54:51)
FIBER016 | 1310 | 0.33 | 0.3 | 0.31 | -0.68 | 0.32 | 43.81 | >44.15 | 0.079
| 1550 | 0.33 | 0.26 | 0.2 | -0.36 | 0.27 | 44.77 | >44.69
(25.03.2008 15:55:11)
FIBER017 | 1310 | 0.54 | 0.3 | 0.52 | -0.68 | 0.53 | 44.03 | >44.53 | 0.067
| 1550 | 0.47 | 0.26 | 0.38 | -0.36 | 0.43 | 44.8 | >45.02
(25.03.2008 15:56:42)
FIBER018 | 1310 | 0.34 | 0.3 | 0.32 | -0.68 | 0.33 | 43.63 | >44.16 | 0.067
| 1550 | 0.26 | 0.26 | 0.18 | -0.36 | 0.22 | 44.51 | >44.65
(25.03.2008 15:57:00)
FIBER019 | 1310 | 0.19 | 0.3 | 0.18 | -0.68 | 0.18 | 43.9 | >43.90 | 0.067
| 1550 | 0.23 | 0.26 | 0.1 | -0.36 | 0.16 | 44.51 | >44.50
(25.03.2008 15:57:18)
FIBER020 | 1310 | 0.31 | 0.3 | 0.31 | -0.68 | 0.31 | >44.18 | >44.15 | 0.068
| 1550 | 0.27 | 0.26 | 0.24 | -0.36 | 0.25 | >45.38 | >44.76
(25.03.2008 15:57:36)
FIBER021 | 1310 | 0.44 | 0.3 | 0.41 | -0.68 | 0.42 | 43.3 | >44.32 | 0.073
| 1550 | 0.28 | 0.26 | 0.27 | -0.36 | 0.27 | 44.45 | >44.81
(25.03.2008 15:59:01)
FIBER022 | 1310 | 0.29 | 0.3 | 0.29 | -0.68 | 0.29 | 42.87 | 43.91 | 0.072
| 1550 | 0.35 | 0.26 | 0.2 | -0.36 | 0.27 | 43.27 | 44.43
(25.03.2008 15:59:22)
FIBER023 | 1310 | 0.13 | 0.3 | 0.16 | -0.68 | 0.15 | 43.49 | 43.68 | 0.072
| 1550 | 0.22 | 0.26 | 0.18 | -0.36 | 0.2 | 43.96 | 43.94
(25.03.2008 15:59:40)
FIBER024 | 1310 | 0.51 | 0.3 | 0.5 | -0.68 | 0.5 | 43.47 | >44.49 | 0.073
| 1550 | 0.38 | 0.26 | 0.35 | -0.36 | 0.36 | 44.24 | >44.95
(25.03.2008 15:59:58)

## TTC-S1D05I
Job Information
Job ID: |  | Date: | 26.03.2008 09:11:54
Reason:
Contractor: |  | Operator A:
Customer: |  | Operator B:
Custom job 1: |  | Custom job 2:
Cable Information
Cable ID: | S1D05I
Location A: |  | Location B:
Cable Mfr.: |  | Cable Type:
Custom cable 1: |  | Custom cable 2:
System Information
| Unit A | Unit B
Unit Name: | FOT-932 | FOT-932
Serial Number: | 417390 | 417389
FasTesT
Fiber ID | Wavelength | Loss | Ref. | Loss | Ref. | Average | ORL | ORL | Length
|  | A->B | A->B | B->A | B->A |  | A | B
| (nm) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB)
(km)
FIBER001 | 1310 | -0.09 | -0.49 | -0.12 | 0.49 | -0.11 | >41.28 | >41.23 | 0.102
| 1550 | 0.03 | -0.27 | 0.01 | 0.43 | 0.02 | >42.28 | >42.97
(26.03.2008 09:12:28)
FIBER002 | 1310 | 0.2 | -0.49 | 0.2 | 0.49 | 0.2 | 14.69 | 13.05 | 0.098
| 1550 | 0.22 | -0.27 | 0.19 | 0.43 | 0.21 | 15.85 | 14.93
(26.03.2008 09:17:33)
FIBER003 | 1310 | 0.96 | -0.49 | 0.93 | 0.49 | 0.95 | 38.53 | 38.38 | 0.099
| 1550 | 0.87 | -0.27 | 0.84 | 0.43 | 0.85 | 39.56 | 39.26
(26.03.2008 09:17:55)
FIBER004 | 1310 | 0.1 | -0.49 | 0.07 | 0.49 | 0.09 | 13.69 | 12.89 | 0.098
| 1550 | 0.37 | -0.27 | 0.4 | 0.43 | 0.39 | 10.85 | 10.39
(26.03.2008 09:18:12)
FIBER005 | 1310 | 0.38 | -0.49 | 0.39 | 0.49 | 0.39 | 20.47 | 19.67 | 0.108
| 1550 | 0.43 | -0.27 | 0.41 | 0.43 | 0.42 | 22.19 | 22.12
(26.03.2008 09:20:01)
FIBER006 | 1310 | 0.18 | -0.49 | 0.35 | 0.49 | 0.26 | 35.26 | 40.01 | 0.107
| 1550 | 0.23 | -0.27 | 0.31 | 0.43 | 0.27 | 36.02 | 40.78
(26.03.2008 09:20:18)
FIBER007 | 1310 | 0.21 | -0.49 | 0.21 | 0.49 | 0.21 | >41.83 | >41.88 | 0.108
| 1550 | 0.23 | -0.27 | 0.24 | 0.43 | 0.23 | >42.65 | >43.41
(26.03.2008 09:20:37)
FIBER008 | 1310 | 0.27 | -0.49 | 0.24 | 0.49 | 0.26 | >41.96 | >41.92 | 0.107
| 1550 | 0.3 | -0.27 | 0.23 | 0.43 | 0.26 | >42.80 | >43.38
(26.03.2008 09:20:54)
FIBER009 | 1310 | 0.4 | -0.49 | 0.37 | 0.49 | 0.39 | 38.18 | 37.23 | 0.105
| 1550 | 0.31 | -0.27 | 0.32 | 0.43 | 0.31 | 29.91 | 31
(26.03.2008 09:22:44)
FIBER010 | 1310 | 0.35 | -0.49 | 0.34 | 0.49 | 0.35 | 23.99 | 23.61 | 0.105
| 1550 | 0.24 | -0.27 | 0.31 | 0.43 | 0.28 | 22.85 | 23.25
(26.03.2008 09:23:05)
FIBER011 | 1310 | 0.34 | -0.49 | 0.3 | 0.49 | 0.32 | 25.34 | 24.82 | 0.105
| 1550 | 0.29 | -0.27 | 0.24 | 0.43 | 0.26 | 26.52 | 26.6
(26.03.2008 09:23:39)
FIBER012 | 1310 | 0.02 | -0.49 | 0.04 | 0.49 | 0.03 | 30.1 | 31.26 | 0.108
| 1550 | 0.02 | -0.27 | 0.03 | 0.43 | 0.02 | 31.56 | 32.89
(26.03.2008 09:23:56)
FIBER013 | 1310 | 0.45 | -0.49 | 0.62 | 0.49 | 0.54 | 29.5 | 31.15 | 0.101
| 1550 | 0.34 | -0.27 | 0.48 | 0.43 | 0.41 | 30.73 | 32.75
(26.03.2008 09:26:05)
FIBER014 | 1310 | 0.39 | -0.49 | 0.37 | 0.49 | 0.38 | 30.66 | 33.82 | 0.1
| 1550 | 0.35 | -0.27 | 0.35 | 0.43 | 0.35 | 32.09 | 35.43
(26.03.2008 09:26:23)
FIBER015 | 1310 | 0.26 | -0.49 | 0.24 | 0.49 | 0.25 | 30.81 | 31.5 | 0.1
| 1550 | 0.28 | -0.27 | 0.2 | 0.43 | 0.24 | 31.96 | 32.96
(26.03.2008 09:26:49)
FIBER016 | 1310 | 0.34 | -0.49 | 0.33 | 0.49 | 0.34 | 30.46 | 31.53 | 0.1
| 1550 | 0.34 | -0.27 | 0.26 | 0.43 | 0.3 | 31.83 | 33.15
(26.03.2008 09:27:06)
FIBER017 | 1310 | 0.93 | -0.49 | 0.87 | 0.49 | 0.9 | 28.66 | 30.74 | 0.089
| 1550 | 0.72 | -0.27 | 0.75 | 0.43 | 0.73 | 29.93 | 32.4
(26.03.2008 09:29:31)
FIBER018 | 1310 | 0.36 | -0.49 | 0.36 | 0.49 | 0.36 | 30.67 | 31.4 | 0.09
| 1550 | 0.37 | -0.27 | 0.3 | 0.43 | 0.34 | 31.78 | 33.06
(26.03.2008 09:29:49)
FIBER019 | 1310 | 0.13 | -0.49 | 0.13 | 0.49 | 0.13 | 31.52 | 32.26 | 0.09
| 1550 | 0.16 | -0.27 | 0.11 | 0.43 | 0.14 | 32.72 | 33.86
(26.03.2008 09:30:07)
FIBER020 | 1310 | 0.72 | -0.49 | 0.69 | 0.49 | 0.71 | 30.95 | 31.35 | 0.089
| 1550 | 0.63 | -0.27 | 0.59 | 0.43 | 0.61 | 31.99 | 33.01
(26.03.2008 09:30:42)
FIBER021 | 1310 | 0.31 | -0.49 | 0.31 | 0.49 | 0.31 | 29.31 | 31.57 | 0.09
| 1550 | 0.41 | -0.27 | 0.33 | 0.43 | 0.37 | 30.58 | 33.2
(26.03.2008 09:32:38)
FIBER022 | 1310 | 0.14 | -0.49 | 0.13 | 0.49 | 0.13 | 30.35 | 31.52 | 0.09
| 1550 | 0.1 | -0.27 | 0.1 | 0.43 | 0.1 | 31.94 | 33.23
(26.03.2008 09:32:57)
FIBER023 | 1310 | 0.29 | -0.49 | 0.31 | 0.49 | 0.3 | 29.5 | 31.21 | 0.09
| 1550 | 0.32 | -0.27 | 0.29 | 0.43 | 0.31 | 30.91 | 32.88
(26.03.2008 09:33:15)
FIBER024 | 1310 | 0.25 | -0.49 | 0.22 | 0.49 | 0.24 | 30.8 | 31.51 | 0.091
| 1550 | 0.2 | -0.27 | 0.23 | 0.43 | 0.22 | 32.39 | 33.26
(26.03.2008 09:35:22)

## TTC-S1D05J
Job Information
Job ID: |  | Date: | 26.03.2008 09:37:19
Reason:
Contractor: |  | Operator A:
Customer: |  | Operator B:
Custom job 1: |  | Custom job 2:
Cable Information
Cable ID: | S1D05J
Location A: |  | Location B:
Cable Mfr.: |  | Cable Type:
Custom cable 1: |  | Custom cable 2:
System Information
| Unit A | Unit B
Unit Name: | FOT-932 | FOT-932
Serial Number: | 417390 | 417389
FasTesT
Fiber ID | Wavelength | Loss | Ref. | Loss | Ref. | Average | ORL | ORL | Length
|  | A->B | A->B | B->A | B->A |  | A | B
| (nm) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB) | (dB)
(km)
FIBER001 | 1310 | 0.16 | -0.49 | 0.2 | 0.49 | 0.18 | 31.72 | 32.14 | 0.099
| 1550 | 0.18 | -0.27 | 0.12 | 0.43 | 0.15 | 33.21 | 33.98
(26.03.2008 09:37:29)
FIBER002 | 1310 | 0.18 | -0.49 | 0.14 | 0.49 | 0.16 | 32.19 | 32.45 | 0.099
| 1550 | 0.11 | -0.27 | 0.16 | 0.43 | 0.14 | 33.65 | 34.19
(26.03.2008 09:37:49)
FIBER003 | 1310 | 0.6 | -0.49 | 0.58 | 0.49 | 0.59 | 34.62 | 37.83 | 0.099
| 1550 | 0.51 | -0.27 | 0.5 | 0.43 | 0.51 | 36.31 | 39.31
(26.03.2008 09:38:07)
FIBER004 | 1310 | 0.19 | -0.49 | 0.18 | 0.49 | 0.19 | 35.13 | 35.31 | 0.1
| 1550 | 0.18 | -0.27 | 0.24 | 0.43 | 0.21 | 36.58 | 36.92
(26.03.2008 09:38:25)
FIBER005 | 1310 | 0.18 | -0.49 | 0.19 | 0.49 | 0.19 | 33.08 | 32.44 | 0.11
| 1550 | 0.13 | -0.27 | 0.21 | 0.43 | 0.17 | 34.35 | 34.27
(26.03.2008 09:39:43)
FIBER006 | 1310 | 0.24 | -0.49 | 0.27 | 0.49 | 0.25 | 34.06 | 34.08 | 0.11
| 1550 | 0.3 | -0.27 | 0.29 | 0.43 | 0.29 | 35.28 | 35.78
(26.03.2008 09:40:03)
FIBER007 | 1310 | 0.33 | -0.49 | 0.32 | 0.49 | 0.33 | 33.39 | 33.37 | 0.11
| 1550 | 0.25 | -0.27 | 0.26 | 0.43 | 0.26 | 34.79 | 35.23
(26.03.2008 09:40:20)
FIBER008 | 1310 | 0.25 | -0.49 | 0.29 | 0.49 | 0.27 | 33.14 | 33.63 | 0.109
| 1550 | 0.2 | -0.27 | 0.2 | 0.43 | 0.2 | 34.7 | 35.46
(26.03.2008 09:40:38)
FIBER009 | 1310 | 0.35 | -0.49 | 0.35 | 0.49 | 0.35 | 32.68 | 33.6 | 0.115
| 1550 | 0.27 | -0.27 | 0.25 | 0.43 | 0.26 | 34.28 | 35.43
(26.03.2008 09:42:36)
FIBER010 | 1310 | 0.5 | -0.49 | 0.53 | 0.49 | 0.52 | 34 | 33.27 | 0.116
| 1550 | 0.4 | -0.27 | 0.43 | 0.43 | 0.41 | 35.38 | 35.11
(26.03.2008 09:42:53)
FIBER011 | 1310 | 0.5 | -0.49 | 0.47 | 0.49 | 0.49 | 33.74 | 34.27 | 0.115
| 1550 | 0.45 | -0.27 | 0.42 | 0.43 | 0.43 | 35.13 | 36.08
(26.03.2008 09:43:10)
FIBER012 | 1310 | 0.66 | -0.49 | 0.68 | 0.49 | 0.67 | 34.77 | 33.42 | 0.116
| 1550 | 0.7 | -0.27 | 0.57 | 0.43 | 0.64 | 35.84 | 35.27
(26.03.2008 09:43:29)
FIBER013 | 1310 | -0.09 | -0.49 | -0.06 | 0.49 | -0.08 | 32.85 | 32.85 | 0.099
| 1550 | 0.03 | -0.27 | -0.03 | 0.43 | 0 | 34.33 | 34.67
(26.03.2008 09:51:59)
FIBER014 | 1310 | 0.34 | -0.49 | 0.38 | 0.49 | 0.36 | 34.33 | 33.39 | 0.099
| 1550 | 0.21 | -0.27 | 0.24 | 0.43 | 0.23 | 35.6 | 35.18
(26.03.2008 09:52:19)
FIBER015 | 1310 | 0.22 | -0.49 | 0.24 | 0.49 | 0.23 | 33.64 | 33.24 | 0.099
| 1550 | 0.11 | -0.27 | 0.21 | 0.43 | 0.16 | 35.14 | 34.98
(26.03.2008 09:52:37)
FIBER016 | 1310 | 0.26 | -0.49 | 0.26 | 0.49 | 0.26 | 33.55 | 33.68 | 0.099
| 1550 | 0.18 | -0.27 | 0.24 | 0.43 | 0.21 | 35.1 | 35.36
(26.03.2008 09:52:56)
FIBER017 | 1310 | 0.14 | -0.49 | 0.18 | 0.49 | 0.16 | 34.04 | 33.84 | 0.093
| 1550 | 0.12 | -0.27 | 0.19 | 0.43 | 0.16 | 35.58 | 35.42
(26.03.2008 09:54:25)
FIBER018 | 1310 | 0.63 | -0.49 | 0.65 | 0.49 | 0.64 | 38.52 | 40.18 | 0.093
| 1550 | 0.54 | -0.27 | 0.43 | 0.43 | 0.49 | 39.57 | 41.57
(26.03.2008 09:57:58)
FIBER019 | 1310 | -0.01 | -0.49 | 0 | 0.49 | 0 | 40.13 | 38.7 | 0.093
| 1550 | 0.12 | -0.27 | 0.07 | 0.43 | 0.09 | 40.86 | 40.01
(26.03.2008 09:58:18)
FIBER020 | 1310 | -0.01 | -0.49 | 0.03 | 0.49 | 0.01 | >41.43 | 39.98 | 0.093
| 1550 | 0.06 | -0.27 | 0.04 | 0.43 | 0.05 | 42.12 | 41.47
(26.03.2008 09:58:36)