## Sample Crate Contents Log
Action |  |  | Contents | [5 empty] | 2-3 |  |  | 4-5
Date | Person | Action Code | Crate Number | CRB Number | PCMB Number | CC Number | CCB Number | MPC Number | DMB | TMB | LINK | DMB | TMB | LINK

## Sample Crate Traveler
Crate |  | Test Results |  |  |  | Test Description |  |  |  |  | Test Notes
Crate Number | CRB Number | Date | Tester's Name | Pass/Fail | Disposition | Location | Stage | Step | Test Name | Test Description
[6 empty] | ISR | 10 | 1 | Gererate Crate Traveler
[6 empty] | ISR | 10 | 2 | Mechanical Assembly
[6 empty] | ISR | 10 | 3 | Install Grounding Bolt
[6 empty] | ISR | 10 | 4 | Install and align BP
[6 empty] | ISR | 11 | 1 | CRB - hoses attached?
[6 empty] | ISR | 11 | 2 | CRB - Tested
[6 empty] | ISR | 12 | 1 | Install CRB
[6 empty] | ISR | 12 | 2 | Attach strain relieves
[6 empty] | ISR | 13 | 1 | Mechanical Inspection
ISR
ISR
ISR
ISR
ISR
ISR
ISR
ISR
ISR
ISR

## Fiber Box Detail
Fiber Box
| 3U top box for connection modules | open for air flow
| 1U bottom tray for slack | NOT open for air flow
Front View |  |  |  | Side View
19" Rack mount
|  | 3U |  |  | 7" or 17.5cm deep
|  | 1U | [5 empty] | 12 1/4 " or 32cm
Front
Top View
The fiber optic cables are in a black plastic tube of about 1U diameter | [7 empty] | tube
this tube enters the back of the 1U tray
this tray Is a sererate part that is bolted to the box
Inside the outer tube are several ~7 smaller tubes
Inside the smaller tubes are the optical fibers
All of the fibers make one coil in the bottom tray and are then
routed to the back of the modules in the 3U box
The modules plug into the front of the 3U box
the fibers from the tube are routed into the back of each module
They are permanently connected inside each module
Each fiber was permanently hand terminated inside the module
The patch fibers are plugged into the front
If it is not mounted inside a rack, the box needs a dust and
mechanical protection cover
[5 empty] | units | inches | cm
[5 empty] | 35.5 | 6.9225 | 17.5831
[5 empty] | 63 | 12.285 | 31.2039
[5 empty] | 10 | 1.95 | 4.953
[5 empty] | 7.5 | 1.4625 | 3.7147

## SX5 Racks
Outline of crates within the test racks used at SX5
Modeled on the two racks for the slice test
| U47 |  | U47
| U46 | Commissioning Rack-1 | U46 | Commissioning Rack-2
| U45 |  | U45
40 | U44 | TTC Crate | U44 | LVPS Crate | xxx
39 | U43 |  | U43
38 | U42 |  | U42
37 | U41 |  | U41
36 | U40 |  | U40
35 | U39 |  | U39
34 | U38 |  | U38
33 | U37 |  | U37
32 | U36 |  | U36
31 | U35 | Fan Tray | U35 | Fan Tray
30 | U34 | Fed Crate | U34 | open
29 | U33 |  | U33 |  | Fan Tray
28 | U32 |  | U32
27 | U31 |  | U31
26 | U30 |  | U30
25 | U29 |  | U29
24 | U28 |  | U28
23 | U27 |  | U27
22 | U26 |  | U26
21 | U25 | Fan Tray | U25 | Fan Tray
20 | U24 | Trigger Crate | U24 | Fan Tray
19 | U23 |  | U23
18 | U22 |  | U22
17 | U21 |  | U21
16 | U20 |  | U20
15 | U19 |  | U19
14 | U18 |  | U18
13 | U17 |  | U17 | PC Crate
12 | U16 |  | U16
11 | U15 | Fan Tray | U15
10 | U14 | Nim Crate | U14
9 | U13 |  | U13
8 | U12 |  | U12
7 | U11 |  | U11
6 | U10 |  | U10
5 | U9 |  | U9
4 | U8 |  | U8
3 | U7 |  | U7
2 | U6 |  | U6
1 | U5 | Fan Tray | U5
| U4 |  | U4
| U3 |  | U3
| U2 |  | U2
| U1 |  | U1

## OSU Tests
Present OSU tests
Prefix | Test # | Name | Boards needed
|  |  | Controller | Target DMB | Target TMB | CCB | MPC | 5 CFEB's | ALCT | LVMB | Notes
OSU | 1 | Prgm VME Prom/FPGA | X | X |  | 0 |  |  |  |  | Test program of VME
OSU | 2 | Prgm MCTRL Prom/FPGA | X | X |  | 0 |  |  |  |  | Test program of MCTRL
OSU | 3 | FIFO | X | X |  | 0 |  |  |  |  | Test operation of FIFO on DMB
OSU | 4 | Volt/Temp | X | X |  | 0 |  | X |  | X | Volt/Temp of each CFEB
OSU | 5 | Low Voltages | X | X |  | 0 |  |  |  | X | ON/OFF for CFEB, ALCT
OSU | 6 | FEB comm | X | X |  | 0 |  | X |  |  | CFEB ID code check
OSU | 7 | Gigabit Comm | X | X |  | 0 |  | X |  |  | high rate data testing
OSU | 8 | Compar Thr | X | X |  | 0 |  | X
OSU | 9 | Pulser Voltage | X | X |  | 0
OSU | 10 | Serial Flash Memory | X | X |  | 0
OSU | 11 | Buckeye Shift FM | X | X |  | 0
OSU | 12 | External Pulser | X | X |  | 0 |  | X
OSU | 13 | Internal Pulser | X | X |  | 0 |  | X
OSU | 14 | Overlap FIFO | X | X |  | 0 |  | X
OSU | 15 | TMB/ALCT Comm | X | X | X | 0 | ? | X | X

## Slot Test
Proposed Peripheral Crate Tests
These are slot by slot tests - moving CSC cables to each DMB/TMB board pair
Test # | Name |  | Boards needed
|  |  | Controller | Target DMB | Target TMB | CCB | MPC | 5 CFEB's | ALCT | LVMB | Notes
| DMB | Volt/Temp | X | X |  |  |  | X
| DMB | Low Voltages | X | X |  |  |  | X |  | X
| DMB | FEB comm | X | X |  |  |  | X
| DMB | Gigabit Comm | X | X |  |  |  | X
| DMB | Compar Thr | X | X |  |  |  | X
| DMB | Pulser Voltage | X | X |  |  |  | X
| DMB | Serial Flash Memory | X | X |  |  |  | X
| DMB | Buckeye Shift FM | X | X |  |  |  | X
| DMB | External Pulser | X | X |  |  |  | X
| DMB | Internal Pulser | X | X |  |  |  | X
| DMB | Overlap FIFO | X | X |  |  |  | X
| DMB | TMB/ALCT Comm | X | ? | X |  |  | ? | X

## B904 Test Software
Software | Personnel |  |  | Test
Yes/No | Programmer | Liason | Hardware Expert
|  |  |  | Series |  |  | Name | Test Details | Status of Software/procedures
N |  |  |  | 21 |  |  | Mechanical test
N |  |  |  | 21 | 1 |  | Mechanical inspection of crate
N |  |  |  | 21 | 1 | 0 | All crate components present, OK and installed correctly
N |  |  |  | 21 | 1 | 1 | Condition of BP and slots – check for bent pins, missing rails ..
N |  |  |  | 21 | 2 | 0 | Install crate in test rack
N |  |  |  | 21 | 2 | 1 | Connect water and turn water on >> leak check
N |  |  |  | 21 | 2 | 2 | Connect LVPS cables and ground bolt to grounding cable. Check continuity
N |  |  |  | 22 |  |  | Power Test
N |  |  |  | 22 | 1 |  | Install PCMB and connect CANBUS cable
Y | Valerii |  |  | 22 | 2 |  | Run ISR PCMB test - | Standared Power ON to Crate Test - Follow special ‘first turn on sequence’ measuring voltages. Software will be an evolution of program first developed for CRB testing at ISR, CRB@ISR. | First version in use. Upgrades underway.
- |  |  |  | 23 |  |  | Install CC board and attach I/O
|  |  |  | 23 | 1 |  | Check that all on-board jumpers are correctly set | This board has XX switches and YY jumpers that must be set correctly TBD
Y | Valerii |  |  | 23 | 2 |  | Power on and check voltages | Standared Power ON to Crate Test - Follow special ‘first turn on sequence’ measuring voltages. Software will be an evolution of program first developed for CRB testing at ISR, CRB@ISR. | First version in use. Upgrades underway.
Y |  |  |  | 23 | 3 |  | Run VME response test
Y |  |  |  | 23 | 4 |  | Run firmware check test
Y |  |  |  | 24 |  |  | Install CCB and attach I/O
|  |  |  | 24 | 1 |  | Check that all on-board jumpers are correctly set | This board has XX switches and YY jumpers that must be set correctly TBD
Y |  |  |  | 24 | 2 |  | Power on and check voltages | Standared Power ON to Crate Test - Follow special ‘first turn on sequence’ measuring voltages. Software will be an evolution of program first developed for CRB testing at ISR, CRB@ISR. | First version in use. Upgrades underway.
Y |  |  |  | 24 | 3 |  | Run VME response test
Y |  |  |  | 24 | 4 |  | Run firmware check test
Y |  |  |  | 24 | 4 |  | Run TTC response test
- |  |  |  | 25 |  |  | Install boards, from right, slot 20/21, working left
N |  |  |  | 25 | 1 |  | Install DMB//TMB/Link in slot pair | Test is done on each DMB//TMB/Link slot pair. For each pair the test CSC chamber is connected
|  |  |  | 25 | 1 | 1 | Check that all on-board jumpers are correctly set | The DMB has XX switches and YY jumpers that must be set correctly. The TMB on-board clock must be disabled. It also has XX switches and YY jumpers that must be set correctly, The LINK board …
Y | Valerii |  |  | 25 | 1 | 2 | Power on and check voltages | Standared Power ON to Crate Test - Follow special ‘first turn on sequence’ measuring voltages. Software will be an evolution of program first developed for CRB testing at ISR, CRB@ISR. | First version in use. Upgrades underway.
Y |  |  |  | 25 | 1 | 3 | OSU01 Prgm VME Prom/FPGA – test program of VME |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 1 | 4 | OSU02 Prgm MCTRL Prom/FPGA – test program of MCTRL |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 1 | 5 | OSU03 FIFO – test operation of FIFO on DMB |  | Have original C program for this test from OSU - Need to adapt to B904
? |  |  |  | 25 | 1 | 6 | Does DMB correctly respond to CCB?
Y |  |  |  | 25 | 1 | 7 | Run VME response test for TMB
Y |  |  |  | 25 | 1 | 8 | Run firmware check test for TMB
? |  |  |  | 25 | 1 | 9 | Does TMB correctly respond to CCB?
- |  |  |  | 25 | 2 |  | Connect CSC cables and FO cables
Y |  |  |  | 25 | 2 | 1 | OSU04 Volt/Temp – CFEB Volt/Temp of each CFEB |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 2 | OSU05 Low Voltages – ON/OFF for CFEB, ALCT |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 3 | OSU06 FEB comm – CFEB ID code check |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 4 | OSU07 Gigabit Comm – high rate data testing |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 5 | OSU08 Compar Thr – |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 6 | OSU09 Pulser Voltage |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 7 | OSU10 Serial Flash Memory |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 8 | OSU11 Bucheye Shift FM |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 9 | OSU12 Exteranl Pulser |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 10 | OSU13 Internal Pulser |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 11 | OSU14 Overlap FIFO |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 12 | OSU15 TMB/ALCT Comm |  | Have original C program for this test from OSU - Need to adapt to B904
Y |  |  |  | 25 | 2 | 13 | Is TMB-DMB communication ok?
Y |  |  |  | 25 | 2 | 14 | Is TMB-ALCT-AFEB communication ok?
- |  |  |  | 26 |  |  | Install MPC
|  |  |  | 26 | 1 |  | Check that all on-board jumpers are correctly set | This board has XX switches and YY jumpers that must be set correctly TBD
Y | Valerii |  |  | 26 | 2 |  | Power on and check voltages | Standared Power ON to Crate Test - Follow special ‘first turn on sequence’ measuring voltages. Software will be an evolution of program first developed for CRB testing at ISR, CRB@ISR. | First version in use. Upgrades underway.
Y |  |  |  | 26 | 3 |  | Run VME response test
Y |  |  |  | 26 | 4 |  | Run firmware check test
? |  |  |  | 26 | 5 |  | Does it respond to CCB?
Y |  |  |  | 26 | 6 |  | Test TMB to MPC communication.
- |  |  |  | 27 |  |  | Intra-Crate Tests
Y |  |  |  | 27 | 1 |  | Test DMB-DDU communication
Y |  |  |  | 27 | 2 |  | Test MPC-SP communication.
Y |  |  |  | 27 | 3 |  | Thrashing Test | Try to run all communication paths simultaneously
N |  |  |  | 28 |  |  | Post Test (passed tests)
N |  |  |  | 28 | 1 |  | De-cable and remove all boards (slots 1-21 and CRBC) to special storage
N |  |  |  | 28 | 2 |  | Store boards in crate at B904? Transport in crate and remove before crate instillation at SX5?
N |  |  |  | 28 | 3 |  | Disconnect crate – clearing all water from CRB and move crate to special storage

## SX5 Test Software
Software | Personnel |  |  | Test
Yes/No | Programmer | Liason | Hardware Expert |  |  |  | B904
|  |  |  | Series |  |  | Number | Name | Test Details | Status of Software/procedures
|  |  |  | 50 |  |  |  | on arrival
|  |  |  | 50 | 1 |  |  | have passed all tests at B904
|  |  |  | 51 |  |  |  | Mechanical test
|  |  |  | 51 | 1 | 1 |  | Mechanical inspection of rack and attachments
|  |  |  | 51 | 1 | 2 |  | Grounding fixtures present
|  |  |  | 51 | 1 | 3 |  | Water to rack
|  |  |  | 51 | 1 | 4 |  | Other lines to rack
|  |  |  | 51 | 1 | 5 |  | mechanical inspection of crate
|  |  |  | 51 | 2 | 1 |  | repeat 21.1.0
|  |  |  | 51 | 2 | 2 |  | repeat 21.1.1
|  |  |  | 51 | 3 |  |  | Install crate in rack
|  |  |  | 51 | 3 | 1 |  | Lift crate to mini-rack and install
|  |  |  | 51 | 3 | 2 |  | Connect water and turn water on >> leak check
|  |  |  | 51 | 3 | 3 |  | Connect Temporary LVPS cables and ground bolt to grounding cable. Check continuity
|  |  |  | 52 |  |  |  | Power Test
|  |  |  | 52 | 1 |  |  | Repeat Test 22 - Power Test at B904
|  |  |  | 53 |  |  |  | Populate crate
|  |  |  | 53 | 1 |  |  | Follow tests 23 through 27 of B904. but now when each board is connected to the skew clear cables they remain attached
|  |  |  | 54 |  |  |  | HV OFF Tests
|  |  |  | 54 | 1 |  |  | Added here are tests beyond those conducted at B904. First run tests that determine boards are cabled to the CSC correctly. Then run calibration tests that begin the determination of timing and other constants for these CSC board combinations.
|  |  |  | 55 |  |  |  | HV ON Tests – These are adopted from the single chamber commissioning tests, (CSC Test) Some or all may be moved to 54 above.
|  |  |  | 55 | 1 |  |  | Slow Control Test (CSC test 9)
|  |  |  | 55 | 2 |  |  | AFEB counting noise (CSC test 11)
|  |  |  | 55 |  |  |  | Test noise of wire groups
|  |  |  | 55 |  |  |  | Check for short circuits
|  |  |  | 55 | 3 |  |  | AFEB Connectivity (CSC test 12)
|  |  |  | 55 |  |  |  | Pulse ALCT to check that all wires are alive
|  |  |  | 55 | 4 |  |  | AFEB Thresholds and Analog Noise (CSC test 13)
|  |  |  | 55 |  |  |  | Measure threshold and noise of each wire group
|  |  |  | 55 | 5 |  |  | CFEB DAQ Path Noise (CSC test 15)
|  |  |  | 55 | 6 |  |  | CFEB Connectivity (CSC test 16)
|  |  |  | 55 | 7 |  |  | Comparator Thresholds and Analog Noise (CSC test 19)
|  |  |  | 56 |  |  |  | Post Test (passed tests)
|  |  |  | 56 | 1 |  |  | Disconnect all ‘Temporary IO’
|  |  |  | 56 | 2 |  |  | Return HV and gas flow as appropriate

## Entire Crate
Proposed Peripheral Crate Tests
These are tests on multiple slots - no cables moves are required
Test # | Name |  | Boards needed
|  |  | Controller | Target DMB | Target TMB | CCB | MPC | 5 CFEB's | ALCT | LVMB | Notes
| DMB | Prgm VME Prom/FPGA | X | X
| DMB | Prgm MCTRL Prom/FPGA | X | X
| DMB | FIFO | X | X
| DMB | Volt/Temp | X | X